## WBS
Specific Electrical Test result
Test of coil, pole or magnet
Project |  |  | : | MQXC
Model |  |  |  | Aperture_2
[5 empty] | Réalisé par : | Francois-Olivier PINCOT | Date: | 04.05.12
| EDMS no.: |  |  |  | Vérifié par : |  | Version no: | 1.06
LHC Project doc: |  |  |  |  | Approuvé par : |  | Date of last modification: | 27.09.13
the
Large
Hadron
Collider
project

## guideline
Procedure general tests electriques
|  |  |  | Prototype | Version no: |  | 1.06
[5 empty] | Date of last modification: |  | 27.09.13
|  |  |  | Project | MQXC |  | Specific Electrical Test result
Test of coil, pole or magnet
|  |  |  | LHC project doc | Page: | 1 | sur | 3
Equipement tests electriques
| A | KEITHLEY 2000 multimeter
| B | Alimentation stabilisee de courant
| C | MEGGER BM21 / S1-1052
| D | Tektronix TDS 3014
| E | LCR Meter 879B BK Precision
| F | HP 4263B LCR meter
| G | SEITZ Impulse tester
| H | EFACEC Heater discharge Power Supply
| I | Labview data transfert software
| J | Press of Young modulus
| K | Tektronix probe P6015A
| L | Connection cable and/or specific rack
Test of single layer, Nb-Ti, after curing splice version - L
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
Lop1 | Coil to copper wedge | Insulation | Coil | Insulation coil to copper wedge | Opt | 500 V | 30 sec |  | C
L1 | Coil to ground | Insulation | Coil | Insulation of the coil in its curing mould to ground | MQC | 500 V | 30 sec | > 1000 MΩ | C
L2 | RDC | Resistance | Coil | Resistance measurement of coil | MQC | 6A | NA | Ref coil +/- 3% | A, B
L2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
L3 | Inductance | Inductance | Coil | Coil inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
L4 | Coil inter-turn discharge | Discharge | Coil | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
Lop2 | Discharge under stress SP | Discharge | Coil | Discharge while putting the straight part in an insulation mould at P=80MPa | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Lop3 | Discharge under stress ends | Discharge | Coil | Discharge while putting the ends in an insulation mould at P=80MPa decreasing to 0 MPa over the ehads | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Test of pole : 2 layers, Nb-Ti, splice version, before Young modulus - BYM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
BYM5 | Coil inter-turn discharge | Discharge | Coil | Discharge test to identify turn to turn problems | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
BYM12 | Q.H. discharge | Discharge | QH | To be applied if Q.H are under coils and already assembled | Opt | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
Test of pole : 2 layers, Nb-Ti, splice version, after Young modulus - AYM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AYM1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Inner layer +outer layer +/- 2% | A, B
AYM2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
AYM2_XXXX | RDC instrumentation | Resistance |  | RDC of each component | MQC | 1A / 4 wires | NA | Compare to previous component measurement +/- 2% | A, B
AYM3 | Inductance | Inductance | Pole | Pole inductance measurement | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref pole] +/- 2% | E
AYM4 | Insulation resistance | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | MQC | 1 kV | 30 sec | > 1000 MΩ | C
AYM5 | Dielectric | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | MQC | 2.4 kV |  | < 5 μA
AYM6 | Coil discharge test | Discharge | Pole | Discharge test to identify turn to turn problems | INV | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AYM7 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
AYM8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
AYM9 | Insulation resistance | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | MQC | 1 kV | 30 sec | < 1000 MΩ | C
AYM10 | Dielectric | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | MQC | 2.4 kV |  | < 5 μA
Test of pole : 1 layer test, Nb-Ti, no splice version - P
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
Pop1 | Pole to copper wedge | Insulation | Pole | Insulation pole to copper wedge | Opt | 500 V | 30 sec |  | C
P1 | Pole to ground | Insulation | Pole | Insulation of the pole in its curing mould to ground | MQC | 500 V | 30 sec | > 1000 MΩ | C
P2 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
P3 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
P3_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
P4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
P5 | Coil inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
Pop2 | Discharge under stress SP | Discharge | Pole | Discharge while putting the straight part in an insulation mould at P=80MPa | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Pop3 | Discharge under stress ends | Discharge | Pole | Discharge while putting the ends in an insulation mould at P=80MPa decreasing to 0 MPa over the ehads | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
P6 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled | MQC | 1 kV | 30 sec | < 1000 MΩ | C
P7 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
P8 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
P9 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled | MQC | 1 kV | 30 sec | < 1000 MΩ | C
P10 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
Test of pole : Nb3Sn, before reaction - PBR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
PBR2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBR3 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PBR4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBR5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before reaction | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PBR6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before reaction | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PBR7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before reaction | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PBR8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
Test of pole : Nb3Sn, after reaction - PAR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled after reaction
PAR2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAR3 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAR4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAR5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled after reaction | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAR6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled after reaction | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAR7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled after reaction | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAR8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled after reaction
Test of pole : Nb3Sn, before impregnation - PBI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PBI2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBI3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PBI4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBI5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PBI6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PBI7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PBI8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAIM1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAIM2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAIM2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAIM3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAIM4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAIM5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAIM6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAIM7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAIM8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
Test of pole : Nb3Sn, after impregnation still out of the mould - PAI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAI2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAI3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAI4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAI5 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
PAI6 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAI7 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
PAI8 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAI9 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAI10 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAI11 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
Test of pole : Nb3Sn, after impregnation still out of the mould 0.1 bar He - PHE
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PHE1 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PHE2 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | 20 V/turn | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | MQC | 40 V/turn
[5 empty] | MQC | 60 V/turn
PHE3 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | Max [(120 V/turn, 5kV/nturn)] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
PHE4 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 250 V | 30 sec | < 1000 MΩ | C
PHE5 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 200 V | 5 min | < 1 μA | C
[5 empty] | INV | 500 V |  | < 2.5 μA
[5 empty] | MQC | 1.7 kV |  | < 10 μA
Test of collared coil before collaring or pre-collaring - BC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
BC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
BC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
BC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
BC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
BC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
BC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
BC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.4 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
BC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.4 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
BC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.4 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
BC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.4 kV | 5 min | < 10 μA | C
All poles to instrumentation
BC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
BC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
BC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.4 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
BC9_XXXX | insulation resistance | Insulation | BSH | instrumentation to ground | MQC | 3.4 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
BC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
BC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
BC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.4 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of collared coil after pre-collaring - APC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
APC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
APC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
APC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
APC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
APC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
APC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
APC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.2 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
APC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.2 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
APC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.2 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
APC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.2 kV | 5 min | < 10 μA | C
All poles to instrumentation
APC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
APC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
APC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.2 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
APC9_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.2 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
APC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
APC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
APC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.2 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of collared coil after collaring - AC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
AC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
AC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
AC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
AC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
AC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
AC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
AC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
AC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
AC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
AC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
AC9_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
AC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
AC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
AC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of the magnet after interconnection - AI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AI1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
AI2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
AI2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AI3 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
AI3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
AI4 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
AI5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
All poles to ground
AI5_XXXX | Insulation resistance | Insulation |  | All Poles + All instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
AI6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
All poles to ground
AI6_XXXX | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
AI7 | Discharge test | Discharge | Pole | All poles in series | INV | Max [(120 V/turn, 5kV/nturn)/2] |  | Previous test | D, G, I, K
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AI8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
AI9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
AI9_XXXX | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to poles
AI10 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
AI11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
AI12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Final Test of the magnet before delivery - FINAL
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
FINAL1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
FINAL2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
FINAL2_MQXC | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
FINAL3 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
FINAL3_MQXC | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
FINAL4 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
FINAL5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
All poles to ground
FINAL5_MQXC | Insulation resistance | Insulation |  | All Poles + All instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
FINAL6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
All poles to ground
FINAL6_MQXC | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
FINAL7 | Discharge test | Discharge | Pole | All poles in series | INV | Max [(120 V/turn, 5kV/nturn)/2] |  | Previous test | D, G, I, K
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
FINAL8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
FINAL9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
FINAL9_MQXC | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to poles
FINAL10 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
FINAL11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
FINAL12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
1) TEST CODE : an identification code for the test
2) TITLE : Test short name
3) TYPE : Type of measure
4) OBJECT : component tested
SH = spot heater, BSH = beam simulation heater, QH = quench heater, TT = temperature sensor
5) SHORT DESCRIPTION: short description of test if necessary
6) PRIORITY of test :
a. Mandatory Q.C. [MQC]: the test shall be performed
b. Optional [Opt] : the test can be performed
c. Investigative [Inv]: the test shall be performed as preparatory test of a mandatory Q.C.
7) TEST CONDITION: i.e. voltage. Two values will be given a set value and a formula to compute the voltage according to previous philosophy above
8) TIME: duration of test if necessary
9) TOLERANCE: values inside which the test is considered to be ok
10) EQUIPMENT: equipment used for the test
the
Large
Hadron
Collider
project
the
Large
Hadron
Collider
project

## Vtest
Summary of tests and of voltages to be applied
| type of magnet | test phase code | Coil/Pole/Aperture Insulation and dielectric to ground | Q.H. Insulation and dielectric to pole and to ground together | Insulation/dielectric instrumentation | Pole Discharge | Q.H. discharge
Series of test | Nb-Ti | L | 200 V | NA | >=1 kV | 120V/turn 5kVmax | NA
|  | BYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | AYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | P | 200 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
| Nb3Sn | PBR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PBI | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAIM | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAI | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PHE | NA | From VQH[i] /10 to VQH[i] /3 | >=1 kV | 120V/turn 5kVmax | 900V 80A
| All magnets | BC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | APC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AI | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | FINAL | Vtaf Table I | VtQHaf Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
L : Test of single layer, Nb-Ti, after curing splice version - L
BYM : Test of pole : 2 layers, Nb-Ti, splice version, before Young modulus – BYM
AYM : Test of pole : 2 layers, Nb-Ti, splice version, after Young modulus – AYM
P : Test of pole : 1 layer test, Nb-Ti, no splice version – P
PBR : Test of pole : Nb3Sn, before reaction – PBR
PAR : Test of pole : Nb3Sn, after reaction – PAR
PBI : Test of pole : Nb3Sn, before impregnation – PBI
PAIM : Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
PAI : Test of pole : Nb3Sn, after impregnation still out of the mould – PAI
PHE : Test of pole : Nb3Sn, after impregnation still out of the mould 0.1 bar He – PHE
BC : Test of collared coil before collaring or pre-collaring – BC
APC : Test of collared coil after pre-collaring – APC
AC : Test of collared coil after collaring – AC
AI : Test of the magnet after interconnection – AI
FINAL : Final Test of the magnet before delivery – FINAL
VQH[i]: Test voltage for the High Voltage test of the Quench Heater to be applied to the ith test
VT[i] : Test voltage for the High Voltage test of the windings to be applied to the ith test

## RDC
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L2 / AYM1 / BC1 / APC1 / AC1 / AI1 / FINAL1
|  |  | Version no: |  | 1.06
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
L2 - RDC Test of single layer, Nb-Ti, after curing splice version
AYM1 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young
BC1 - RDC Mesure de resistance de chaque pole et des V-taps avant pre-collaring
APC1 - RDC Mesure de resistance de chaque pole et des V-taps apres pre-collaring
AC1 - RDC Mesure de resistance de chaque pole et des V-taps apres collaring
AI1 - RDC Mesure de resistance de chaque pole et des V-taps apres interconnection
FINAL1 - RDC Mesure de resistance de chaque pole et des V-taps avant livraison
Appareil utilise : |  | KEITHLEY 2000 multimeter
Alimentation stabilisee de courant
courant aux bornes de chaque pole [A] |  |  |  | 4
| Nom du test |  | L2
| Tests effectués le : |  | 2012-08-14
| Nom de l'opérateur : |  | FO Pincot
| Tamb [°C] |  | 25.1
| Humidité [%] |  | 45.3
| Coil name |  | Mesurée
|  |  | Tension | Resistance |  |  | Resistance
|  |  | [mV] | [mW] |  |  | [mW]
| CERN_06_inner |  | 294 | 73.5 |  | inner average | 72.8875
| CERN_07_inner |  | 287.6 | 71.9 |  | std deviation | 0.89454
| CERN_08_inner |  | 295.1 | 73.775 |  | outer average | 99.0906
| CERN_09_inner |  | 289.5 | 72.375 |  | std deviation | 1.6565
| CERN_06_outer |  | 400.6 | 100.15
| CERN_07_outer |  | 403.4 | 100.85
| CERN_08_outer |  | 390.15 | 97.5375
| CERN_09_outer |  | 391.3 | 97.825
courant aux bornes de chaque pole [A] |  |  |  | 6
|  | Nom du test |  |  | AYM1 - bf pre-collaring |  | BC1 - af pre-collaring |  | APC1 - af pre-collaring |  | AC1 - af collaring
|  | Tests effectués le : |  |  | 2012-10-17 | [5 empty] | 2012-10-31
|  | Nom de l'opérateur : |  |  | FOP | [5 empty] | FOP
|  | Tamb [°C] |  |  | 21.5 | [5 empty] | 22.4
|  | Humidité [%] |  |  | 35.4 | [5 empty] | 24.1
Coil 7
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner | inner coil 7 | 976.45 | 162.7417 |  | 0 |  | 0 | 957.45 | 159.575 | Previous measured | +/- 1%
|  | U+ |  | EEI71 | 568 | 94.6667 |  | 0 |  | 0 | 553.38 | 92.23 | Previous measured | +/- 1%
|  | U+ |  | EEI72 | 568 | 94.6667 |  | 0 |  | 0 | 553.38 | 92.23 | Previous measured | +/- 1%
|  | U+ | outer | EEO71 | 567.6 | 94.6 |  | 0 |  | 0 | 552.99 | 92.165 | Previous measured | +/- 1%
|  | U+ |  | EEO72 | 567.6 | 94.6 |  | 0 |  | 0 | 553 | 92.1667 | Previous measured | +/- 1%
| I- | U- |  | outer coil 7 | 12 | 2 |  | 0 |  | 0 | 0.18 | 0.03 | Previous measured | +/- 1%
|  | Nom du test |  |  | AYM1 - bf pre-collaring |  | BC1 - af pre-collaring |  | APC1 - af pre-collaring |  | AC1 - af collaring
|  | Tests effectués le : |  |  | 2012-10-17 | [5 empty] | 2012-10-31
|  | Nom de l'opérateur : |  |  | FOP | [5 empty] | FOP
|  | Tamb [°C] |  |  | 21.5 | [5 empty] | 22.4
|  | Humidité [%] |  |  | 35.4 | [5 empty] | 24.1
Coil 3
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner | inner coil 3 | 981.5 | 163.5833 |  | 0 |  | 0 | 965.52 | 160.92 | Previous measured | +/- 1%
|  | U+ |  | EEHI31 | 581.78 | 96.9633 |  | 0 |  | 0 | 566.08 | 94.3467 | Previous measured | +/- 1%
|  | U+ |  | EEHI32 | 581.76 | 96.96 |  | 0 |  | 0 | 566.1 | 94.35 | Previous measured | +/- 1%
|  | U+ |  | EEJI31 | 574.4 | 95.7333 |  | 0 |  | 0 | 558.31 | 93.0517 | Previous measured | +/- 1%
|  | U+ |  | EEJI32 | 574.4 | 95.7333 |  | 0 |  | 0 | 558.31 | 93.0517 | Previous measured | +/- 1%
|  | U+ |  | EEI31 | 573.8 | 95.6333 |  | 0 |  | 0 | 557.73 | 92.955 | Previous measured | +/- 1%
|  | U+ |  | EEI32 | 573.8 | 95.6333 |  | 0 |  | 0 | 557.74 | 92.9567 | Previous measured | +/- 1%
|  | U+ | outer | EEO31 | 573.5 | 95.5833 |  | 0 |  | 0 | 557.32 | 92.8867 | Previous measured | +/- 1%
|  | U+ |  | EEO32 | 573.5 | 95.5833 |  | 0 |  | 0 | 557.33 | 92.8883 | Previous measured | +/- 1%
| I- | U- |  | outer coil 3 | 2.41 | 0.40167 |  | 0 |  | 0 | 0.17 | 0.028333 | Previous measured | +/- 1%
|  | Nom du test |  |  | AYM1 - bf pre-collaring |  | BC1 - af pre-collaring |  | APC1 - af pre-collaring |  | AC1 - af collaring
|  | Tests effectués le : |  |  | 2012-10-17 | [5 empty] | 2012-10-31
|  | Nom de l'opérateur : |  |  | FOP | [5 empty] | FOP
|  | Tamb [°C] |  |  | 21.5 | [5 empty] | 22.4
|  | Humidité [%] |  |  | 35.4 | [5 empty] | 24.1
Coil 8
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner | inner coil 8 | 971.05 | 161.8417 |  | 0 |  | 0 | 957.3 | 159.55 | Previous measured | +/- 1%
|  | U+ |  | EEI81 | 565.09 | 94.1817 |  | 0 |  | 0 | 553.15 | 92.1917 | Previous measured | +/- 1%
|  | U+ |  | EEI82 | 565.08 | 94.18 |  | 0 |  | 0 | 553.15 | 92.1917 | Previous measured | +/- 1%
|  | U+ | outer | EEO81 | 564.66 | 94.11 |  | 0 |  | 0 | 552.74 | 92.1233 | Previous measured | +/- 1%
|  | U+ |  | EEO82 | 564.65 | 94.1083 |  | 0 |  | 0 | 552.74 | 92.1233 | Previous measured | +/- 1%
| I- | U- |  | outer coil 8 | 10.4 | 1.7333 |  | 0 |  | 0 | 0.28 | 0.046667 | Previous measured | +/- 1%
|  | Nom du test |  |  | AYM1 - bf pre-collaring |  | BC1 - af pre-collaring |  | APC1 - af pre-collaring |  | AC1 - af collaring
|  | Tests effectués le : |  |  | 2012-10-17 | [5 empty] | 2012-10-31
|  | Nom de l'opérateur : |  |  | FOP | [5 empty] | FOP
|  | Tamb [°C] |  |  | 21.5 | [5 empty] | 22.4
|  | Humidité [%] |  |  | 35.4 | [5 empty] | 24.1
Coil 9
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner | inner coil 9 | 981.16 | 163.5267 |  | 0 |  | 0 | 959.5 | 159.9167 | Previous measured | +/- 1%
|  | U+ |  | EEHI91 | 584.6 | 97.4333 |  | 0 |  | 0 | 563.62 | 93.9367 | Previous measured | +/- 1%
|  | U+ |  | EEHI92 | 584.1 | 97.35 |  | 0 |  | 0 | 563.63 | 93.9383 | Previous measured | +/- 1%
|  | U+ |  | EEJI91 | 573 | 95.5 |  | 0 |  | 0 | 555.84 | 92.64 | Previous measured | +/- 1%
|  | U+ |  | EEJI92 | 573 | 95.5 |  | 0 |  | 0 | 555.83 | 92.6383 | Previous measured | +/- 1%
|  | U+ |  | EEI91 | 572.1 | 95.35 |  | 0 |  | 0 | 555.26 | 92.5433 | Previous measured | +/- 1%
|  | U+ |  | EEI92 | 572.1 | 95.35 |  | 0 |  | 0 | 555.27 | 92.545 | Previous measured | +/- 1%
|  | U+ | outer | EEO91 | 571.4 | 95.2333 |  | 0 |  | 0 | 554.84 | 92.4733 | Previous measured | +/- 1%
|  | U+ |  | EEO92 | 571.6 | 95.2667 |  | 0 |  | 0 | 554.86 | 92.4767 | Previous measured | +/- 1%
| I- | U- |  | outer coil 9 | 2.07 | 0.345 |  | 0 |  | 0 | 0.12 | 0.02 | Previous measured | +/- 1%
courant aux bornes de chaque pole [A] |  |  |  | 6
|  | Nom du test |  |  | AI1 |  | FINAL1 |  | FINAL Vtaps protec |  | FINAL1 Modified
|  | Tests effectués le : |  |  |  |  | 21/12/2012 |  | 15/01/2013 |  | 2013-09-24
|  | Nom de l'opérateur : |  |  |  |  | FOP |  | FOP |  | FOP
|  | Tamb [°C] |  |  |  |  | 20.3 |  | 20.7 |  | 23.2
|  | Humidité [%] |  |  |  |  | 29.5 |  | 19.1 |  | 40.5
All coils in serie
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner coil9 | inner coil 9 |  | 0 | 3816 | 636 | 3771 | 628.5 | 3788 | 631.3333 | Previous measured | +/- 1%
|  | U+ |  | EEHI91 |  | 0 | 3409 | 568.1667 | 3373 | 562.1667 | 3391 | 565.1667 | Previous measured | +/- 1%
|  | U+ |  | EEHI92 |  | 0 | 3409 | 568.1667 | 3373 | 562.1667 | 3383.6 | 563.9333 | Previous measured | +/- 1%
|  | U+ |  | EEJI91 |  | 0 | 3402 | 567 | 3365 | 560.8333 | 3384.1 | 564.0167 | Previous measured | +/- 1%
|  | U+ |  | EEJI92 |  | 0 | 3402 | 567 | 3365 | 560.8333 | 3384.1 | 564.0167 | Previous measured | +/- 1%
|  | U+ |  | EEI91 |  | 0 | 3402 | 567 | 3365 | 560.8333 | 3383.6 | 563.9333 | Previous measured | +/- 1%
|  | U+ |  | EEI92 |  | 0 | 3402 | 567 | 3365 | 560.8333 | 3391.9 | 565.3167 | Previous measured | +/- 1%
|  | U+ | outer coil9 | EEO91 |  | 0 | 3401 | 566.8333 | 3364 | 560.6667 | 3383 | 563.8333 | Previous measured | +/- 1%
|  | U+ |  | EEO92 |  | 0 | 3401 | 566.8333 | 3364 | 560.6667 | 3383 | 563.8333 | Previous measured | +/- 1%
|  | U+ | splice 9-3 | EESO91 |  | 0 | 2858 | 476.3333 | 2822 | 470.3333 | 2837.4 | 472.9 | Previous measured | +/- 1%
|  | U+ |  | EESO92 |  | 0 | 2858 | 476.3333 | 2822 | 470.3333 | 2837.3 | 472.8833 | Previous measured | +/- 1%
|  | U+ |  | EESI31 |  | 0 | 2858 | 476.3333 | 2822 | 470.3333 | 2837.3 | 472.8833 | Previous measured | +/- 1%
|  | U+ |  | EESI32 |  | 0 | 2858 | 476.3333 | 2822 | 470.3333 | 2837.3 | 472.8833 | Previous measured | +/- 1%
|  | U+ |  | EESI33 | didn't exist | #VALUE! | didn't exist | #VALUE! | didn't exist | #VALUE! | 2839.1 | 473.1833 | Previous measured | +/- 1%
|  | U+ |  | EESI34 | didn't exist | #VALUE! | didn't exist | #VALUE! | didn't exist | #VALUE! | 2839.1 | 473.1833 | Previous measured | +/- 1%
|  | U+ | inner coil3 | EEHI31 |  | 0 | 2470 | 411.6667 | 2435 | 405.8333 | 2446.8 | 407.8 | Previous measured | +/- 1%
|  | U+ |  | EEHI32 |  | 0 | 2470 | 411.6667 | 2435 | 405.8333 | 2446.8 | 407.8 | Previous measured | +/- 1%
|  | U+ |  | EEJI31 |  | 0 | 2461 | 410.1667 | 2427 | 404.5 | 2439 | 406.5 | Previous measured | +/- 1%
|  | U+ |  | EEJI32 |  | 0 | 2461 | 410.1667 | 2427 | 404.5 | 2438.9 | 406.4833 | Previous measured | +/- 1%
|  | U+ |  | EEI31 |  | 0 | 2460 | 410 | 2426 | 404.3333 | 2438.4 | 406.4 | Previous measured | +/- 1%
|  | U+ |  | EEI32 |  | 0 | 2460 | 410 | 2426 | 404.3333 | 2438.3 | 406.3833 | Previous measured | +/- 1%
|  | U+ | outer coil3 | EEO31 |  | 0 | 2460 | 410 | 2426 | 404.3333 | 2437.9 | 406.3167 | Previous measured | +/- 1%
|  | U+ |  | EEO32 |  | 0 | 2460 | 410 | 2426 | 404.3333 | 2437.9 | 406.3167 | Previous measured | +/- 1%
|  | U+ | splice 3-7 | EESO31 |  | 0 | HS | #VALUE! | 1885 | 314.1667 | 1893.4 | 315.5667 | Previous measured | +/- 1%
|  | U+ |  | EESO32 |  | 0 | 1918 | 319.6667 | 1885 | 314.1667 | 1893.3 | 315.55 | Previous measured | +/- 1%
|  | U+ |  | Lead out3out7 |  | 0 | didn't exist | #VALUE! | didn't exist | #VALUE! | 1894.3 | 315.7167 | Previous measured | +/- 1%
|  | U+ |  | EESO71 |  | 0 | 1918 | 319.6667 | HS | #VALUE! | HS | #VALUE! | Previous measured | +/- 1%
|  | U+ |  | EESO72 |  | 0 | 1918 | 319.6667 | 1885 | 314.1667 | 1893.3 | 315.55 | Previous measured | +/- 1%
|  | U+ | outer coil7 | EEO71 |  | 0 | 1373 | 228.8333 | 1341 | 223.5 | 1346.1 | 224.35 | Previous measured | +/- 1%
|  | U+ |  | EEO72 |  | 0 | 1372 | 228.6667 | 1341 | 223.5 | 1346 | 224.3333 | Previous measured | +/- 1%
|  | U+ | inner coil7 | EEI71 |  | 0 | 1372 | 228.6667 | 1341 | 223.5 | 1346 | 224.3333 | Previous measured | +/- 1%
|  | U+ |  | EEI72 |  | 0 | 1372 | 228.6667 | 1341 | 223.5 | 1346 | 224.3333 | Previous measured | +/- 1%
|  | U+ | splice 7-8 | EESI71 |  | 0 | 974.9 | 162.4833 | 945 | 157.5 | 946.8 | 157.8 | Previous measured | +/- 1%
|  | U+ |  | EESI72 |  | 0 | 974.9 | 162.4833 | 945.8 | 157.6333 | 946.83 | 157.805 | Previous measured | +/- 1%
|  | U+ |  | Lead in7out8 |  | 0 | 974.9 | 162.4833 | 945.7 | 157.6167 | 947.74 | 157.9567 | Previous measured | +/- 1%
|  | U+ |  | EESO81 |  | 0 | 974.9 | 162.4833 | 944.7 | 157.45 | 946.78 | 157.7967 | Previous measured | +/- 1%
|  | U+ |  | EESO82 |  | 0 | 974.9 | 162.4833 | 944.8 | 157.4667 | 946.79 | 157.7983 | Previous measured | +/- 1%
|  | U+ | outer coil8 | EEO81 |  | 0 | 443.6 | 73.9333 | 403 | 67.1667 | 401.31 | 66.885 | Previous measured | +/- 1%
|  | U+ |  | EEO82 |  | 0 | 443.6 | 73.9333 | 403 | 67.1667 | 401.3 | 66.8833 | Previous measured | +/- 1%
|  | U+ | inner coil8 | EEI81 |  | 0 | 432.7 | 72.1167 | 402.6 | 67.1 | 400.89 | 66.815 | Previous measured | +/- 1%
|  | U+ |  | EEI82 |  | 0 | 432.7 | 72.1167 | 402.6 | 67.1 | 400.88 | 66.8133 | Previous measured | +/- 1%
|  | U+ |  | inner coil 8 |  | 0 | 48.5 | 8.0833 | 0.2 | 0.033333 | 0.07 | 0.011667 | Previous measured | +/- 1%

## heater RDC
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES SPECIFIQUES
Test of coil, pole or magnet
Test code AYM2-8 / BC2-8-11 / AC2-8-11 / APC2-8-11 / AI2-8-11 / FINAL2-8-11
|  |  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
AYM2 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young
AYM8 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young et apres test decharge
BC 2 - RDC - Mesure initiale de resistance de chaque Quench Heater (QH) - avant pre-collaring
BC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - avant pre-collaring
BC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge
APC 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres pre-collaring
APC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres pre-collaring
APC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres pre-collaring
AC 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres collaring
AC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres collaring
AC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres collaring
AI 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres interconnection
AI 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres interconnection
AI 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres interconnection
FINAL 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - avant livraison
FINAL 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - avant livraison
FINAL 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - avant livraison
Appareil utilise : |  |  | KEITHLEY 2000 multimeter
Alimentation stabilisee de courant
courant aux bornes de chaque QH [A] |  |  |  | 1
|  | Nom du test |  |  | AYM2 - initial |  | AYM8 - af discharge
|  | Tests effectués le : |  |  | 2012-10-17
|  | Nom de l'opérateur : |  |  | FOP
|  | Tamb [°C] |  |  | 21.8
|  | Humidité [%] |  |  | 39.7
R[dc] Dipole Voltage Taps |  |  |  | nom | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [V] | [W] | [V] | [W] | [W]
Quench Heater
| I+ |  | U+ | YTI7L | 7.419 | 7.419 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO7L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI7R | 7.381 | 7.381 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO7R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI3L | 7.4 | 7.4 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO3L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI3R | 7.407 | 7.407 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO3R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI8L | 7.341 | 7.341 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO8L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L1 | 7.788 | 7.788 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L1 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L2 | 7.788 | 7.788 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L2 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI8R | 7.386 | 7.386 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO8R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R1 | 7.837 | 7.837 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R1 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R2 | 7.836 | 7.836 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R2 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI9L | HS | #VALUE! |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO9L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI9R | 7.383 | 7.383 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO9R |  |  |  |  | Previous measured | +/- 1%
Spot Heater
| I+ |  | U+ | SHI3 | 6.834 | 17.085 | 0.4 | A
| I- |  | U- | SHO3
| I+ |  | U+ | SHI9 | 6.724 | 16.81 | 0.4 | A
| I- |  | U- | SHO9
Temperature sensor (4 wires measure)
|  |  |  | TTI3 |  |  | no connector
|  |  |  | TTI9 |  |  | no connector
|  | Nom du test |  |  | AC2 - initial |  | AC11 - af discharge
|  | Tests effectués le : |  |  | 2012-10-31 |  | 2012-11-02
|  | Nom de l'opérateur : |  |  | FOP |  | FOP
|  | Tamb [°C] |  |  | 23 |  | 22.5
|  | Humidité [%] |  |  | 24.6 |  | 28.9
R[dc] Dipole Voltage Taps |  |  |  | nom | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [V] | [W] | [V] | [W] | [W]
Quench Heater
| I+ |  | U+ | YTI7L | 7.41 | 7.41 | 7.41 | 7.41 | Previous measured | +/- 1%
| I- |  | U- | YTO7L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI7R | 7.44 | 7.44 | 7.37 | 7.37 | Previous measured | +/- 1%
| I- |  | U- | YTO7R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI3L | 7.39 | 7.39 | 7.39 | 7.39 | Previous measured | +/- 1%
| I- |  | U- | YTO3L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI3R | 7.39 | 7.39 | 7.4 | 7.4 | Previous measured | +/- 1%
| I- |  | U- | YTO3R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI8L | 7.33 | 7.33 | 7.34 | 7.34 | Previous measured | +/- 1%
| I- |  | U- | YTO8L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L1 | 7.78 | 7.78 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L1 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L2 | 7.78 | 7.78 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L2 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI8R | 7.37 | 7.37 | 7.38 | 7.38 | Previous measured | +/- 1%
| I- |  | U- | YTO8R |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R1 | 7.83 | 7.83 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R1 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R2 | 7.82 | 7.82 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R2 |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI9L | HS | #VALUE! | HS | #VALUE! | Previous measured | +/- 1%
| I- |  | U- | YTO9L |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | YTI9R | 7.38 | 7.38 | 7.38 | 7.38 | Previous measured | +/- 1%
| I- |  | U- | YTO9R |  |  |  |  | Previous measured | +/- 1%
Beam Simulation Heater
| I+ |  | U+ | BSI73 | 1.72 | 1.72
| I- |  | U- | BSO73
| I+ |  | U+ | EBSI73 | 1.98 | 1.98
| I- |  | U- | EBSO73
| I+ |  | U+ | BSI73 | 0.94 | 0.94
| I- |  | U- | BSM73
| I+ |  | U+ | EBSI73 | 1.156 | 1.156
| I- |  | U- | EBSM73
| I+ |  | U+ | BSI38 | 1.74 | 1.74
| I- |  | U- | BSO38
| I+ |  | U+ | EBSI38 | 2.02 | 2.02
| I- |  | U- | EBSO38
| I+ |  | U+ | BSI89 | 1.78 | 1.78
| I- |  | U- | BSO89
| I+ |  | U+ | EBSI89 | 2.08 | 2.08
| I- |  | U- | EBSO89
| I+ |  | U+ | BSI97 | 1.73 | 1.73
| I- |  | U- | BSO97
| I+ |  | U+ | EBSI97 | 2.01 | 2.01
| I- |  | U- | EBSO97
Spot Heater
| I+ |  | U+ | SHI3 | 6.94 | 17.35 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SHO3 |  |  | Previous measured | +/- 1%
| I+ |  | U+ | SHI9 | 7.17 | 17.925 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SHO9 |  |  | Previous measured | +/- 1%
Temperature sensor (4 wires measure)
|  |  |  | TTI3 |  | HS
|  |  |  | TTI9 |  | 25250
|  |  |  | TTCI |  | 940.28
|  |  |  | TTCO |  | 944.75
|  | Nom du test |  |  | FINAL2 - initial |  | FINAL8 - bf discharge |  | FINAL11 - af discharge |  | FINAL2 modified - initial |  | FINAL11 modified - af discharge
|  | Tests effectués le : |  |  | 2012-12-21 |  |  |  | 2013-01-07 |  | 2013-09-25 |  | 2013-09-25
|  | Nom de l'opérateur : |  |  | FOP |  |  |  | FOP |  | FOP |  | FOP
|  | Tamb [°C] |  |  | 21.5 |  |  |  | 21.2 |  | 23.2 |  | 23.2
|  | Humidité [%] |  |  | 28.3 |  |  |  | 26.3 |  | 40.5 |  | 40.5
R[dc] Dipole Voltage Taps |  |  |  | nom | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [V] | [W] | [V] | [W] | [V] | [W] | [V] | [W] | [V] | [W] | [W]
Quench Heater
| I+ |  | U+ | YTI7L | 7.402 | 7.402 |  | 0 | 7.391 | 7.391 | 7.426 | 7.426 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO7L | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI7R | 7.364 | 7.364 |  | 0 | 7.352 | 7.352 | 7.386 | 7.386 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO7R | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI3L | 7.381 | 7.381 |  | 0 | 7.37 | 7.37 | 7.404 | 7.404 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO3L | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI3R | 7.39 | 7.39 |  | 0 | 7.379 | 7.379 | 7.412 | 7.412 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO3R | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI8L | 7.327 | 7.327 |  | 0 | 7.391 | 7.391 | 7.348 | 7.348 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO8L | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L1 | 7.774 | 7.774 |  | 0 |  | 0 | 7.797 | 7.797 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L1 | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L2 | 7.771 | 7.771 |  | 0 |  | 0 | 7.796 | 7.796 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L2 | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI8R | 7.372 | 7.372 |  | 0 | 7.361 | 7.361 | 7.394 | 7.394 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO8R | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R1 | 7.817 | 7.817 |  | 0 |  | 0 | 7.84 | 7.84 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R1 | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R2 | 7.817 | 7.817 |  | 0 |  | 0 | 7.843 | 7.843 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R2 | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI9L | HS | #VALUE! |  | 0 | HS | #VALUE! | HS | #VALUE! | HS | #VALUE! | Previous measured | +/- 1%
| I- |  | U- | YTO9L | [10 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI9R | 7.369 | 7.369 |  | 0 | 7.358 | 7.358 | 7.391 | 7.391 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | YTO9R | [10 empty] | Previous measured | +/- 1%
Beam Simulation Heater
| I+ |  | U+ | BSI73 | 1.724 | 1.724 | 1.722 | 1.722 |  |  | 1.723 | 1.723 |  | 0
| I- |  | U- | BSO73
| I+ |  | U+ | EBSI73 | 1.971 | 1.971 | 1.968 | 1.968 |  |  | 1.973 | 1.973 |  | 0
| I- |  | U- | EBSO73
| I+ |  | U+ | BSI73 | 0.8743 | 0.8743 | 0.906 | 0.906 |  |  | 0.908 | 0.908 |  | 0
| I- |  | U- | BSM73
| I+ |  | U+ | EBSI73 | 1.1523 | 1.1523 | 1.426 | 1.426 |  |  | 1.434 | 1.434 |  | 0
| I- |  | U- | EBSM73
| I+ |  | U+ | BSI38 | 1.738 | 1.738 |  |  |  |  | 1.737 | 1.737
| I- |  | U- | BSO38
| I+ |  | U+ | EBSI38 | 2.011 | 2.011 |  |  |  |  | 2.012 | 2.012
| I- |  | U- | EBSO38
| I+ |  | U+ | BSI89 | 1.778 | 1.778 |  |  |  |  | 1.779 | 1.779
| I- |  | U- | BSO89
| I+ |  | U+ | EBSI89 | 2.082 | 2.082 |  |  |  |  | 2.082 | 2.082
| I- |  | U- | EBSO89
| I+ |  | U+ | BSI97 | 1.732 | 1.732 |  |  |  |  | 1.732 | 1.732
| I- |  | U- | BSO97
| I+ |  | U+ | EBSI97 | 2.008 | 2.008 |  |  |  |  | 2.008 | 2.008
| I- |  | U- | EBSO97
Spot Heater
| I+ |  | U+ | SH3 | 6.935 | 17.3375 | Previous measured | +/- 1% | 0.4 | A | 6.93 | 17.325 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SH3 |  |  | Previous measured | +/- 1% |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | SH9 | 7.095 | 17.7375 | Previous measured | +/- 1% | 0.4 | A | HS | #VALUE! | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SH9 |  |  | Previous measured | +/- 1% |  |  |  |  | Previous measured | +/- 1%
| I+ |  | U+ | SHSI3 | didn't exist | #VALUE! | Previous measured | +/- 1% | 0.4 | A | 6.283 | 15.7075 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SHSI3 |  |  | Previous measured | +/- 1% |  |  |  |  | Previous measured | +/- 1%
Temperature sensor (4 wires measure)
|  |  |  | TTI3 |  | HS | [5 empty] | HS
|  |  |  | TTI9 |  | HS | [5 empty] | HS
|  |  |  | TTS39 |  | 912.15 | [5 empty] | 911.04
|  |  |  | TTSI3 |  | 998.32 | [5 empty] | 996.75
|  |  |  | TTCI |  | 944.82
|  |  |  | TTCO |  | 940.3
|  |  |  | TTL |  | 983.95 | [5 empty] | 983.79

## Capacitance
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code BC3 / APC3 / AC3 / AI3 / FINAL3
|  |  | Version no: | 1.06
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
BC3 - Capacitance - Mesure de la capacite des poles avant pre-collaring
APC3 - Capacitance - Mesure de la capacite des poles apres pre-collaring
AC3 - Capacitance - Mesure de la capacite des poles apres collaring
AI3 - Capacitance - Mesure de la capacite des poles apres interconnection
FINAL3 - Capacitance - Mesure de la capacite des poles avant livraison
Appareil utilise : |  | LCR Meter 879B BK Precision
Mesures realisees a une frequence de (Hz): |  |  | 1000
| Nom du test |  | BC3 - bf pre-collaring |  | APC3 - af pre-collaring | AC3 - af collaring
| Tests effectués le : |  |  |  |  | 2012-11-01
| Nom de l'opérateur : |  |  |  |  | FOP
| Tamb [°C] |  |  |  |  | 21.9
| Humidité [%] |  |  |  |  | 28.6
Capacitance |  |  | Mesurée |  | Mesurée | Mesurée
|  |  | [nF] |  | [nF] | [nF]
QH7L ---> Pole 7 | [5 empty] | 4.04
QH7R ---> Pole 7 | [5 empty] | 4.08
QH3L ---> Pole 3 | [5 empty] | 4.11
QH3R ---> Pole 3 | [5 empty] | 4.04
QH8L ---> Pole 8 | [5 empty] | 4.34
QH8R ---> Pole 8 | [5 empty] | 4.29
QH9L ---> Pole 9 | [5 empty] | 4.11
QH9R ---> Pole 9 | [5 empty] | 4.06
Pole 7 ---> Ground | [5 empty] | 7.47
Pole 3 ---> Ground | [5 empty] | 7.51
Pole 8 ---> Ground | [5 empty] | 7.58
Pole 9 ---> Ground | [5 empty] | 7.54
Pole 7 ---> Pole 3 | [5 empty] | 4.99
Pole 7 ---> Pole 8 | [5 empty] | 4.22
Pole 7 ---> Pole 9 | [5 empty] | 4.92
Pole 3 ---> Pole 8 | [5 empty] | 4.99
Pole 3 ---> Pole 9 | [5 empty] | 4.27
Pole 8 ---> Pole 9 | [5 empty] | 4.95
All poles ---> Ground | [5 empty] | 18.86
| Nom du test |  | FINAL3 - Final |  | FINAL3 - Final modified
| Tests effectués le : |  | 2012-12-21 |  | 2013-09-24
| Nom de l'opérateur : |  | FOP |  | FOP
| Tamb [°C] |  | 22.1 |  | 23.6
| Humidité [%] |  | 29.4 |  | 42.9
Capacitance |  |  | Mesurée |  | Mesurée
|  |  | [nF] |  | [nF]
QH7L ---> all poles |  |  | 3.91 |  | 3.97
QH7R ---> all poles |  |  | 3.95 |  | 4
QH3L ---> all poles |  |  | 3.98 |  | 4.05
QH3R ---> all poles |  |  | 3.93 |  | 3.99
QH8L ---> all poles |  |  | 4.18 |  | 4.2
QH8R ---> all poles |  |  | 4.13 |  | 4.15
QH9L ---> all poles |  |  | 4.03 |  | 4.08
QH9R ---> all poles |  |  | 3.98 |  | 4.05
All poles ---> Ground |  |  | 17.47 |  | 18.38
Tolerances :
Compare single pole among them and to previous assembly

## Inductance
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L3 / AYM3 / BC4 / APC4 / AC4 / AI4 / FINAL4
|  |  |  | Version no: |  | 1.06
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
L3 -Test of single layer, Nb-Ti, after curing splice version
AYM3 - Mesure de l'inductance des poles apres mesure module d'Young
BC4 - Mesure de l'inductance des poles avant pre-collaring
APC4 - Mesure de l'inductance des poles apres pre-collaring
AC4 - Mesure de l'inductance des poles apres collaring
AI4 - Mesure de l'inductance des poles apres interconnection
FINAL4 - Mesure de l'inductance des poles avant livraison
Appareil utilise : |  | BK PRECISION LCR meter 879B
Tests effectués le : |  | 2012-08-13
Nom de l'opérateur : |  | FO Pincot
Tamb [°C] |  | 26
Humidité [%] |  | 48.9
L3 | after Young modulus | Inductance [µH]
|  | Mesure réalisée |  | CERN_06_inner | CERN_07_inner | CERN_08_inner | CERN_09_inner | CERN_06_outer | CERN_07_outer | CERN_08_outer | CERN_09_outer
|  | Fréquence | 100 Hz | 322 | 309 | 305 | 308 | 488 | 495 | 493 | 490
|  |  | 1000 HZ | 319.2 | 306.4 | 298.5 | 305.7 | 485.1 | 492.1 | 485 | 486.1
|  |  | 1 0000 Hz | 310.3 | 297.3 | 286.2 | 296.8 | 472.4 | 477.9 | 468.5 | 472.6
Tests effectués le : |  | 2012-10-17
Nom de l'opérateur : |  | FOP
Tamb [°C] |  | 21.8
Humidité [%] |  | 40.2
AYM3 | after Young modulus | Inductance [mH]
|  | Mesure réalisée |  | pole 7 | pole 3 | pole 8 | pole 9
|  | Fréquence | 100 Hz | 1.319 | 1.336 | 1.325 | 1.319
|  |  | 1000 HZ | 1.292 | 1.302 | 1.3 | 1.292
|  |  | 1 0000 Hz | 1.229 | 1.235 | 1.24 | 1.228
Tests effectués le :
Nom de l'opérateur :
Tamb [°C]
Humidité [%]
BC4 | before pre-collaring | Inductance [mH]
|  | Mesure réalisée |  | pole 7 | pole 3 | pole 8 | pole 9 | all poles in serie
|  | Fréquence | 100 Hz
1000 HZ
1 0000 Hz
Tests effectués le :
Nom de l'opérateur :
Tamb [°C]
Humidité [%]
APC4 | after pre-collaring | Inductance [mH]
|  | Mesure réalisée |  | pole 7 | pole 3 | pole 8 | pole 9 | all poles in serie
|  | Fréquence | 100 Hz
1000 HZ
1 0000 Hz
Tests effectués le : |  | 2012-10-31
Nom de l'opérateur : |  | FOP
Tamb [°C] |  | 22.9
Humidité [%] |  | 25.4
AC4 | after collaring | Inductance [mH]
|  | Mesure réalisée |  | pole 7 | pole 3 | pole 8 | pole 9 | all poles in serie
|  | Fréquence | 100 Hz | 1,56 | 1,56 | 1,57 | 1,57 | 8,67
|  |  | 1000 HZ | 1,19 | 1,20 | 1,20 | 1,20 | 7,10
|  |  | 1 0000 Hz | 0,68 | 0,681 | 0,679 | 0,678 | 3,82
Tests effectués le : |  | 2012-12-21 |  | 2013-09-24
Nom de l'opérateur : |  | FOP |  | FOP
Tamb [°C] |  | 22.7 |  | 23.6
Humidité [%] |  | 27.6 |  | 42.9
Inductance [mH]
Test Code |  | FINAL |  | FINAL modified
Mesure réalisée |  | all poles in serie |  | all poles in serie
Fréquence | 100 Hz | 8.26 |  | 8.26
| 1000 HZ | 7.1 |  | 7.09
| 1 0000 Hz | 4.31 |  | 4.32
Tolerances :
Compare single pole among them and to previous assembly

## Insulation
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM 4-9 / BC5 / APC5 / AC5 / AI5 / FINAL5
|  |  | Version no: |  | 1.06
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
AYM4 - Mesure de resistance d'isolation des poles apres module d'Young
AYM9 - Mesure de resistance d'isolation des poles apres module d'Young apres test decharge
BC5 - Mesure de resistance d'isolation des poless avant pre-collaring
APC5 -Mesure de resistance d'isolation des poles apres pre-collaring
AC5 - Mesure de resistance d'isolation des poles apres collaring
AI5 - Mesure de resistance d'isolation des poles apres interconnection
FINAL5 - Mesure de resistance d'isolation des poles avant livraison
Appareil utilise : |  | MEGGER BM21
| Nom du test |  | AYM 4 |  |  | AYM9
| Tests effectués le : |  | 2012-10-17 |  |  | 2012-10-18
| Nom de l'opérateur : |  | FOP |  |  | FOP
| Tamb [°C] |  | 22.3 |  |  | 22
| Humidité [%] |  | 42.4 |  |  | 37.4
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [GW]
Pole 7 ---> QH7L + QH7R |  |  | 1000 | 30 | 505 | 1000 | 30 | 458 | >1
Pole 3 ---> QH3L + QH3R |  |  | 1000 | 30 | 98.7 | 1000 | 30 | 149 | >1
Pole 8 ---> QH8L + QH8R |  |  | 1000 | 30 | 83.8 | 1000 | 30 | 57.3 | >1
Pole 9 ---> QH9L + QH9R |  |  | 1000 | 30 | 93.1 | 1000 | 30 | 101 | >1
| Nom du test |  | BC5 - bf pre-collaring |  |  | APC5 - af pre-collaring |  |  | AC5 - af collaring |  |  | AC5 - mandrin take out
| Tests effectués le : |  | 2012-10-24 | [5 empty] | 2012-11-01 |  |  | 2012-11-05
| Nom de l'opérateur : |  | N.Bourcey | [5 empty] | FOP |  |  | FOP
| Tamb [°C] |  | 23 | [5 empty] | 22.1 |  |  | 22.3
| Humidité [%] |  | 45 | [5 empty] | 28.1 |  |  | 34.1
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [GW]
All poles + all QHs ---> Ground |  |  | 1400 | 30 | 21.6 | 1200 | 30 |  | 1000 | 30 | 17.5 | 1000 | 30 | 32.9 | >1
all poles ---> all QH |  |  | 1400 | 30 | 230 | 1200 | 30 |  | 1000 | 30 | 28 | 1000 | 30 | 27.4 | >1
Pole 7 ---> Ground |  |  | 1400 | 30 | 46.1 | 1200 | 30 |  | 1000 | 30 | 45.5 | 1000 | 30 |  | >1
Pole 3 ---> Ground |  |  | 1400 | 30 | 49.7 | 1200 | 30 |  | 1000 | 30 | 41.9 | 1000 | 30 |  | >1
Pole 8 ---> Ground |  |  | 1400 | 30 | 42.7 | 1200 | 30 |  | 1000 | 30 | 50.3 | 1000 | 30 |  | >1
Pole 9 ---> Ground |  |  | 1400 | 30 | 51.6 | 1200 | 30 |  | 1000 | 30 | 44.7 | 1000 | 30 |  | >1
Pole 7 ---> Pole 3 |  |  | 1400 | 30 | 75.8 | 1200 | 30 |  | 1000 | 30 | 52.2 | 1000 | 30 | 62.6 | >1
Pole 7 ---> Pole 8 |  |  | 1400 | 30 | 65.6 | 1200 | 30 |  | 1000 | 30 | 85.4 | 1000 | 30 | 77 | >1
Pole 7 ---> Pole 9 |  |  | 1400 | 30 | 84.8 | 1200 | 30 |  | 1000 | 30 | 107 | 1000 | 30 | 68.8 | >1
Pole 3 ---> Pole 8 |  |  | 1400 | 30 | 60.6 | 1200 | 30 |  | 1000 | 30 | 49 | 1000 | 30 | 56.4 | >1
Pole 3 ---> Pole 9 |  |  | 1400 | 30 | 77.7 | 1200 | 30 |  | 1000 | 30 | 35.8 | 1000 | 30 | 51.9 | >1
Pole 8 ---> Pole 9 |  |  | 1400 | 30 | 75.6 | 1200 | 30 |  | 1000 | 30 | 80.9 | 1000 | 30 | 60.5 | >1
| Nom du test |  | FINAL5 - Final |  |  | FINAL5 - Vtaps protec |  |  | FINAL5 - modified
| Tests effectués le : |  | 2012-12-21 |  |  | 2013-01-15 |  |  | 2013-09-24
| Nom de l'opérateur : |  | FOP |  |  | FOP |  |  | FOP
| Tamb [°C] |  | 21.8 |  |  | 20.7 |  |  | 23.6
| Humidité [%] |  | 29.6 |  |  | 19.1 |  |  | 42.9
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [GW]
All poles + all QHs ---> Ground |  |  | 1000 | 30 | 47.7 | 1000 | 30 |  | 1000 | 30 | 17.9 | >1
all poles ---> all QHs |  |  | 1000 | 30 | 76.5 | 1000 | 30 |  | 1000 | 30 | 22.5 | >1
all poles ---> Ground |  |  | 1000 | 30 | 44.5 | 1000 | 30 | 41.7 | 1000 | 30 | 16.1 | >1

## Dielectric
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM5-10 / BC6 / APC6 / AC6 / AI6 / FINAL6
|  |  | Version no: |  | 1.06
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
AYM5 - Mesure du courant de fuite dans le dielectrique apres module d'Young
AYM10 - Mesure du courant de fuite dans le dielectriques apres module d'Young et apres test decharge
BC6 - Mesure du courant de fuite dans le dielectrique avant pre-collaring
APC6 - Mesure du courant de fuite dans le dielectriques apres pre-collaring
AC6 - Mesure du courant de fuite dans le dielectrique apres collaring
AI6 - Mesure du courant de fuite dans le dielectrique apres interconnection
FINAL6 - Mesure du courant de fuite dans le dielectrique avant livraison
Appareil utilise : |  | MEGGER BM21
| Nom du test |  | AYM5 |  |  | AYM10
| Tests effectués le : |  | 2012-10-17 |  |  | 2012-10-18
| Nom de l'opérateur : |  | FOP |  |  | FOP
| Tamb [°C] |  | 22.3 |  |  | 22
| Humidité [%] |  | 42.4 |  |  | 37.4
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [min] | [nA] | [V] | [min] | [nA] | [µA]
Pole 7 ---> QH7L + QH7R |  |  | 1000 | 5 | 0.61 | 1000 | 5 | 0.72 | < 2
Pole 7 ---> QH7L + QH7R |  |  | 2400 | 5 | 2.01 | 2400 | 5 | 1.93 | < 5
Pole 3 ---> QH3L + QH3R |  |  | 1000 | 5 | 2.73 | 1000 | 5 | 1.81 | < 2
Pole 3 ---> QH3L + QH3R |  |  | 2400 | 5 | 7.61 | 2400 | 5 | 5.35 | < 5
Pole 8 ---> QH8L + QH8R |  |  | 1000 | 5 | 2.26 | 1000 | 5 | 3.91 | < 2
Pole 8 ---> QH8L + QH8R |  |  | 2400 | 5 | 12.8 | 2400 | 5 | 12.3 | < 5
Pole 9 ---> QH9L + QH9R |  |  | 1000 | 5 | 4.29 | 1000 | 5 | 2.71 | < 2
Pole 9 ---> QH9L + QH9R |  |  | 2400 | 5 | 11.5 | 2400 | 5 | 8.67 | < 5
| Nom du test |  | BC6 - bf pre-collaring |  |  | APC6 - af pre-collaring |  |  | AC6 - af collaring
| Tests effectués le : |  | 24&25/10/2012 | [5 empty] | 2012-11-01
| Nom de l'opérateur : |  | N.Bourcey | [5 empty] | FOP
| Tamb [°C] |  | 23 / 21.9 | [5 empty] | 22.1
| Humidité [%] |  | 45 / 47.5 | [5 empty] | 28.1
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [min] | [µA] | [V] | [min] | [µA] | [V] | [min] | [µA] | [µA]
All poles + all QHs ---> Ground |  |  | 3000 | 5 | 0.065 | 2800 | 5 |  | 2600 | 5 | 0.044 | < 10
all poles ---> all QH |  |  | 3000 | 5 | 0.042 | 2800 | 5 |  | 2600 | 5 | 0.018 | < 10
Pole 7 ---> Ground |  |  | 3000 | 5 | 0.023 | 2800 | 5 |  | 2600 | 5 | 0.016 | < 10
Pole 3 ---> Ground |  |  | 3000 | 5 | 0.014 | 2800 | 5 |  | 2600 | 5 | 0.016 | < 10
Pole 8 ---> Ground |  |  | 3000 | 5 | 0.023 | 2800 | 5 |  | 2600 | 5 | 0.013 | < 10
Pole 9 ---> Ground |  |  | 3000 | 5 | 0.017 | 2800 | 5 |  | 2600 | 5 | 0.015 | < 10
Pole 7 ---> Pole 3 |  |  | 3000 | 5 | 0.008 | 2800 | 5 |  | 2600 | 5 | 0.01 | < 10
Pole 7 ---> Pole 8 |  |  | 3000 | 5 | 0.012 | 2800 | 5 |  | 2600 | 5 | 0.007 | < 10
Pole 7 ---> Pole 9 |  |  | 3000 | 5 | 0.007 | 2800 | 5 |  | 2600 | 5 | 0.008 | < 10
Pole 3 ---> Pole 8 |  |  | 3000 | 5 | 0.008 | 2800 | 5 |  | 2600 | 5 | 0.01 | < 10
Pole 3 ---> Pole 9 |  |  | 3000 | 5 | 0.007 | 2800 | 5 |  | 2600 | 5 | 0.01 | < 10
Pole 8 ---> Pole 9 |  |  | 3000 | 5 | 0.009 | 2800 | 5 |  | 2600 | 5 | 0.007 | < 10
| Nom du test |  | FINAL6 - Final |  |  | FINAL6 - Vtaps protec |  |  | FINAL6 - modified
| Tests effectués le : |  | 2012-12-21 |  |  | 2013-01-15 |  |  | 2013-09-24
| Nom de l'opérateur : |  | FOP |  |  | FOP |  |  | FOP
| Tamb [°C] |  | 21.8 |  |  | 20.7 |  |  | 23.6
| Humidité [%] |  | 29.6 |  |  | 19.1 |  |  | 42.9
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [min] | [µA] | [V] | [min] | [µA] | [V] | [min] | [µA] | [µA]
All poles + all QHs ---> Ground |  |  | 2400 | 5 | 0.0149 | 2400 | 5 |  | 2400 | 5 | 0.063 | < 10
all poles ---> all QHs |  |  | 2400 | 5 | 0.007 | 2400 | 5 |  | 2400 | 5 | 0.01 | < 10
All poles ---> Ground |  |  | 2400 | 5 | 0.0142 | 2400 | 5 | 0.0187 | 2400 | 5 | 0.066 | < 10

## Discharge
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L4 / AYM6 / BC7 / APC7 / AC7 / AI7 / FINAL7
|  |  |  | Version no: |  | 1.06
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
L4 - Test of single layer, Nb-Ti, after curing splice version
AYM6 - Discharge test - Mesure du temps de decharge des poles apres module d'Young
BC7 - Discharge test - Mesure du temps de decharge des poles avant pre-collaring
APC7 - Discharge test - Mesure du temps de decharge des poles apres pre-collaring
AC7 - Discharge test - Mesure du temps de decharge des poles apres collaring
AI7 - Discharge test - Mesure du temps de decharge des poles apres interconnection
FINAL7 - Discharge test - Mesure du temps de decharge des poles avant livraison
Appareil utilise : |  |  | SEITZ Impulse tester
Maximum 120V/turn and 5kV/nturn
Nom du test |  | L4 |  |  | inner coil : 17 turns |  | measure at 1.7kV and 2.04kV
[5 empty] | outer coil : 19 turns |  | measure at 1.9kV and 2.3kV
Tests effectués le : |  | 2012-08-14
Nom de l'opérateur : |  | FO Pincot
Tamb [°C] |  | 25.1
Humidité [%] |  | 45.3
Graphique
Nom du test |  | AYM6 |  | BC7 |  | APC7 |  | AC7 | AI7 | FINAL7 | FINAL7 | 36 turns by pole | measure at 3.6kV and 4.3kV
|  | af Young modulus |  | bf pre-collaring |  | af pre-collaring |  | af collaring | af interconnection | Final | Final modified | after interconnection measure at 5.0kV
Tests effectués le : |  | 2012-10-17 | [5 empty] | 2012-11-02 |  | 2013-01-07 | 2013-09-24
Nom de l'opérateur : |  | FOP | [5 empty] | FOP |  | FOP | FOP
Tamb [°C] |  | 22.3 | [5 empty] | 22.2 |  | 20.7 | 24.6
Humidité [%] |  | 42.4 | [5 empty] | 28.9 |  | 26.8 | 46.8
Constante de temps Ƭ |  | Ƭ Mesurée |  | Ƭ Mesurée |  | Ƭ Mesurée |  | Ƭ Mesurée | Ƭ Mesurée | Ƭ Mesurée | Ƭ Mesurée
|  | [µs] |  | [µs] |  | [µs] |  | [µs] | [µs] | [µs] | [µs]
Pole 7 | [8 empty] | All poles | All poles | All poles
Pole 3
Pole 8
Pole 9
Graphique

## Heater insulation
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES SPECIFIQUES
Test of coil, pole or magnet
Test code BC9-12 / APC9-12 / AC9-12 / AI9-12 / FINAL 9-12
|  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
BC9 - Mesure de resistance d'isolation des heaters avant pre-collaring
BC12 - Mesure de resistance d'isolation des heaters apres test decharge avant pre-collaring
APC9 - Mesure de resistance d'isolation des heaters apres pre-collaring
ACP12 - Mesure de resistance d'isolation des heaters apres test decharge apres pre-collaring
AC9 - Mesure de resistance d'isolation des heaters apres collaring
AC12 - Mesure de resistance d'isolation des heaters apres test decharge apres collaring
AI9 - Mesure de resistance d'isolation des heaters apres interconnection
AI12 - Mesure de resistance d'isolation des heaters apres test decharge apres interconnection
FINAL9 - Mesure de resistance d'isolation des heaters avant livraison
FINAL12 - Mesure de resistance d'isolation des heaters apres test decharge avant livraison
Appareil utilise : |  | MEGGER BM21
| Tests effectués le : |  | 2012-10-25 | [13 empty] | 2012-11-01 |  |  | 2012-11-02 |  |  | 2012-11-05
| Nom de l'opérateur : |  | N.Bourcey | [13 empty] | FOP |  |  | FOP |  |  | FOP
| Tamb [°C] |  | 21.9 | [13 empty] | 23.2 |  |  | 22.5 |  |  | 22.3
| Humidité [%] |  | 47.5 | [13 empty] | 28.2 |  |  | 28.9 |  |  | 34.1
| Nom du test |  | BC9 - before discharge test |  |  |  |  | BC12 - af discharge test |  |  | APC9 - bf discharge test |  |  | APC12 - af discharge test |  |  | AC9 - bf discharge test |  |  | AC12 - af discharge test |  |  | AC9 - mandrin take out
Insulation resistance |  |  | U[test] | duration | Mesurée | nominale |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée
|  |  | [V] | [s] | [GW] | [GW] |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW]
All QHs --> Ground |  |  | 3400 | 60 | 5.8 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 54.4 | 3000 | 60 | 44.6 | 3000 | 60 | 87.4
All BSHs --> Ground |  |  | 3400 | 60 | 43.9 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 47.5 | 3000 | 60 | 32.8 | 3000 | 60 | 70.2
All SHs --> Ground |  |  | 3400 | 60 | 785 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 1420 | 3000 | 60 | 1380 | 3000 | 60
QH 7L --> Pole 7 |  |  | 3400 | 60 | 179 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 517 | 3000 | 60 | 222 | 3000 | 60
QH 7R --> Pole 7 |  |  | 3400 | 60 | 272 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 560 | 3000 | 60 | 261 | 3000 | 60
QH 3L --> Pole 3 |  |  | 3400 | 60 | 272 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 248 | 3000 | 60 | 237 | 3000 | 60
QH 3R --> Pole 3 |  |  | 3400 | 60 | 256 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 233 | 3000 | 60 | 246 | 3000 | 60
QH 8L --> Pole 8 |  |  | 3400 | 60 | 230 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 368 | 3000 | 60 | 340 | 3000 | 60
QH 8R --> Pole 8 |  |  | 3400 | 60 | 235 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 394 | 3000 | 60 | 367 | 3000 | 60
QH 9L --> Pole 9 |  |  | 3400 | 60 | 134 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 325 | 3000 | 60 | 223 | 3000 | 60
QH 9R --> Pole 9 |  |  | 3400 | 60 | 190 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 452 | 3000 | 60 | 401 | 3000 | 60
BSH7-3 --> Pole 7 |  |  | 3400 | 60 | 105 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 139 | 3000 | 60 | 114 | 3000 | 60
BSH7-3 --> Pole 3 |  |  | 3400 | 60 | 210 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 117 | 3000 | 60 | 114 | 3000 | 60
BSH3-8 --> Pole 3 |  |  | 3400 | 60 | 142 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 107 | 3000 | 60 | 102 | 3000 | 60
BSH3-8 --> Pole 8 |  |  | 3400 | 60 | 108 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 147 | 3000 | 60 | 107 | 3000 | 60
BSH8-9 --> Pole 8 |  |  | 3400 | 60 | 209 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 116 | 3000 | 60 | 103 | 3000 | 60
BSH8-9 --> Pole 9 |  |  | 3400 | 60 | 132 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 104 | 3000 | 60 | 109 | 3000 | 60
BSH9-7 --> Pole 9 |  |  | 3400 | 60 | 97.3 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 109 | 3000 | 60 | 90.5 | 3000 | 60
BSH9-7 --> Pole 7 |  |  | 3400 | 60 | 187 | > | 1 | 3400 | 60 |  | 3200 | 60 |  | 3200 | 60 |  | 3000 | 60 | 169 | 3000 | 60 | 171 | 3000 | 60
SH3 --> Pole 3 |  |  | 1000 | 60 | 1120 | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 3000 | 60 | 1910 | 1000 | 60 | 2790 | 3000 | 60
SH9 --> Pole 9 |  |  | 1000 | 60 | 1180 | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 3000 | 60 | 2470 | 1000 | 60 | 3000 | 3000 | 60
TT3 --> Ground |  |  | 1000 | 60 | 1180 | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | 60.3 |  |  |  | 1000 | 60 | 90.6
TT9 --> Ground |  |  | 1000 | 60 | 1180 | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | 70.7 |  |  |  | 1000 | 60 | 80
TT3 --> Pole 3 |  |  | 1000 | 60 |  | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | short |  |  |  | 1000 | 60
TT9 --> Pole 9 |  |  | 1000 | 60 |  | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | short |  |  |  | 1000 | 60
TTCI --> Ground |  |  | 1000 | 60 |  | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | > 3000 |  |  |  | 1000 | 60 | 1630
TTCO --> Ground |  |  | 1000 | 60 |  | > | 1 | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 |  | 1000 | 60 | 1670 |  |  |  | 1000 | 60 | 1810
| Tests effectués le : | [9 empty] | 2013-01-07 |  |  | 2013-09-26
| Nom de l'opérateur : | [9 empty] | FOP |  |  | FOP
| Tamb [°C] | [9 empty] | 21.2 |  |  | 23.8
| Humidité [%] | [9 empty] | 26.3 |  |  | 45.5
| Nom du test |  | AI9 - af interconnection |  |  |  |  | AI12 - af interconnect. |  |  | FINAL12 - Final |  |  | FINAL12 - Final modified
Insulation resistance |  |  | U[test] | duration | Mesurée |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | U[test] | duration | Mesurée
|  |  | [V] | [s] | [GW] |  |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [V] | [s] | [GW]
All QHs --> Ground |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 126 | 3000 | 60 | 46
All BSHs --> Ground |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 102 | 3000 | 60 | 37.7
All SHs --> Ground |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 2500 | 60 | 2340 | 2500 | 60 | 1500
QH 9L --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 833 | 3000 | 60 | 408
QH 9R --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 986 | 3000 | 60 | 487
QH 8L --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 703 | 3000 | 60 | 347
QH 8R --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 788 | 3000 | 60 | 353
QH 3L --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 722 | 3000 | 60 | 287
QH 3R --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 737 | 3000 | 60 | 299
QH 7L --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 828 | 3000 | 60 | 459
QH 7R --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 999 | 3000 | 60 | 577
BSH7-3 -->all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 115 | 3000 | 60 | 70.6
BSH3-8 --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 129 | 3000 | 60 | 71.9
BSH8-9 --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 132 | 3000 | 60 | 76
BSH9-7 --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 169 | 3000 | 60 | 97.8
SH3 --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 4330 | 2500 | 60 | 2100
SH9 --> all poles |  |  | 3000 | 60 |  |  |  | 3000 | 60 |  | 3000 | 60 | 4240 | 2500 | 60 | 3020
SHSI3 --> all poles | [13 empty] | 2500 | 60 | 5010
TT3 --> Ground |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | HS | 1000 | 60 | HS
TT9 --> Ground |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | HS | 1000 | 60 | HS
TT3 --> All poles |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | HS | 1000 | 60 | HS
TT9 --> All poles |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | HS | 1000 | 60 | HS
TTS39 --> Ground |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 2360 | 1000 | 60 | 1010
TTSI3 --> Ground |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 3000 | 1000 | 60 | 1880
TTS39 --> All poles |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 2040 | 1000 | 60 | 1100
TTSI3 --> All poles |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 3000 | 1000 | 60 | 1830
Poles --> Ground - TTCI/CO/L connect gnd |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 78.1 | 1000 | 60
Instru --> Ground - TTCI/CO/L connect gnd |  |  | 1000 | 60 |  |  |  | 1000 | 60 |  | 1000 | 60 | 80.8 | 1000 | 60

## QH discharge
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM7 / BC10 / APC10 / AC10 / AI10 / FINAL10
|  |  |  | Version no: | 1.06
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
AYM7 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres module d'Young
BC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH avant pre-collaring
APC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres pre-collaring
AC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres collaring
AI10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres interconnection
FINAL10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH avant livraison
Appareil utilise : |  |  | EFACEC Heater discharge Power Supply
Tektronix TDS 3014
Labview data transfert software
condition test [V] |  |  |  | 450
heater discharge power capacity [F] |  |  |  | 0.0071
|  | Ƭ Theorique |  | Ƭ=RC
Nom du test |  | AYM7 |  | BC10 | APC10 | AC10 | AI10 | FINAL10
|  | bf pre-collaring |  | bf pre-collaring | af pre-collaring | af collaring | af interconnection | Final
Tests effectués le : |  | 2012-10-18 |  |  |  | 2012-11-02 |  | 2013-01-07
Nom de l'opérateur : |  | FOP |  |  |  | FOP |  | FOP
Tamb [°C] |  | 21.5 |  |  |  | 22.2 |  | 20.3
Humidité [%] |  | 38.1 |  |  |  | 28.9 |  | 27.1
Constante de temps Ƭ |  | Ƭ Mesurée |  | Ƭ Mesurée | Ƭ Mesurée | Ƭ Mesurée | Ƭ Mesurée | Ƭ Mesurée | Ƭ Theorique
|  | [ms] |  | [ms] | [ms] | [ms] | [ms] | [ms] | [ms]
QH7L | [8 empty] | 52.54
QH7R | [8 empty] | 52.54
QH3L | [8 empty] | 52.54
QH3R | [8 empty] | 52.54
QH8L | [8 empty] | 52.54
QH8R | [8 empty] | 52.54
QH9L |  | HS |  |  |  | HS |  | different response | 52.54
QH9R | [8 empty] | 52.54
Graphique