## Guideline
Procedure general tests electriques
|  |  |  | Prototype | Version no: |  | 1
[5 empty] | Date of last modification |  | 28.08.2013
|  |  |  | Project | RMC |  | Specific Electrical Test result
[5 empty] | RMC#Copper |  | Test of coil, pole or magnet
|  |  |  | LHC project doc | Page: | 1 | sur | 8
Electrical test Equipment
| A | DMM as KEITHLEY 2000 multimeter
| B | Stabilized DC current supply as Agilent 6631B
| C | MEGGER BM21 / MEGGER S1-1052
| D | Oscilloscope as Tektronix TDS 3014
| E | LCR Meter 879B BK Precision / equipment F
| F | LCR meter as HP 4263B LCR meter
| G | SEITZ Impulse tester
| H | EFACEC Heater discharge Power Supply
| I | Labview data transfert software
| J | Press of Young modulus
| K | Tektronix probe P6015A
| L | Connection cable and/or specific rack
Test of pole : Nb3Sn, before reaction - PBR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
PBR3_XXXX | RDC | Resistance |  | Resistance measurement of instrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBR4 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 50 V | 60 sec | > 1000 MΩ | C
PBR5 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBR6 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled and completed before reaction | Inv | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PBR7 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled and completed before reaction | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
PBR8 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before reaction | MQC | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PBR9 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
Test of pole : Nb3Sn, after reaction - PAR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled after reaction
PAR2_XXXX | RDC | Resistance |  | Resistance measurement of instrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAR3 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 50 V | 60 sec | > 1000 MΩ | C
PAR4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAR5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled and completed before reaction | Inv | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PAR6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled and completed before reaction | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
PAR7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled and completed before reaction | MQC | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PAR8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are between coils and already assembled and completed before reaction
Test of pole : Nb3Sn, before impregnation - PBI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PBI2_XXXX | RDC | Resistance |  | Resistance measurement of instrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBI3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 50 V | 60 sec | > 1000 MΩ | C
PBI4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBI5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled and completed before impregnation | Inv | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PBI6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled and completed before impregnation | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
PBI7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled and completed before impregnation | MQC | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PBI8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are between coils and already assembled and completed before impregnation
Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAIM1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAIM2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAIM2_XXXX | RDC | Resistance |  | Resistance measurement of instrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAIM3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | VT[i] | 30 sec | > 1000 MΩ | C
PAIM4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAIM5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PAIM6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
PAIM7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | VQ[i] | 30 sec. | >1000 MΩ; < 2 µΩ | C
5 min
PAIM8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
Test of pole : Nb3Sn, after impregnation out of the mould - PAI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAI2_XXXX | RDC | Resistance |  | Resistance measurement of instrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAI3 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAI4 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | Inv | 30 V/turn | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | 60 V/ turn
[5 empty] | MQC | 120 V/turn
PAI5 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | VQ[i]/4 | 30 sec. | >1000 MΩ; < 2 μA | C
[5 empty] | INV | VQ[i]/2 | 5 min | >1000 MΩ; < 5 μA
[5 empty] | MQC | VQ[i] |  | >500 MΩ; < 10 μA
PAI6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
PAI7 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
PAI8 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | VQ[i]/4 | 30 sec. | >1000 MΩ; < 2 μA | C
[5 empty] | INV | VQ[i]/2 | 5 min | >1000 MΩ; < 5 μA
[5 empty] | MQC | VQ[i] |  | >500 MΩ; < 10 μA
Test of pole : Nb3Sn, after impregnation 0.1 bar He - PHE
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PHE1 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PHE5 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | VT[i]/10 | 30 sec. | >1000 MΩ; < 1 μA | C
[5 empty] | INV | VT[i]/5 | 5 min | >1000 MΩ; < 2.5 μA
[5 empty] | MQC | VT[i]/3 |  | >500 MΩ; < 10 μA
PHE2 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | 10 V/turn | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | MQC | 20 V/turn
[5 empty] | MQC | 40 V/turn
PHE3 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | 60 V/turn | 10 puls | Ref coil τ +/- 1% | D, G, I, K
PHE5 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | VT[i]/10 | 30 sec. | >1000 MΩ; < 1 μA | C
[5 empty] | INV | VT[i]/5 | 5 min | >1000 MΩ; < 2.5 μA
[5 empty] | MQC | VT[i]/3 |  | >500 MΩ; < 10 μA
Test of collared coil before collaring or pre-collaring - BC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
BC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6:00 | NA | Previous measured value +/- 1% | A, B
BC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1:00 | NA | Previous measured value +/- 1% | A, B
BC3_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
BC4 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
BC5_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
BC6 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
BC7 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to all Q.H. |  |  | 5 min
|  |  |  | Single pole to each other pole |  | VT[i]
BC8_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to instrumentation |  |  | 5 min
BC9 | Discharge test | Discharge | Pole | Single poles | INV | 30 V/turn |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | 60 V/turn
[5 empty] | MQC | 100 V/turn
BC10 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1:00 |  | Previous tests | A, B
BC11 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
BC12_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 1000 V | 30 sec. | > 1000 MΩ | C
instrumentation to its poles
BC13 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
BC14 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1:00 |  | BC8 | A, B
BC15 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
Test of collared coil after pre-collaring - APC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
APC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6:00 | NA | Previous measured value +/- 1% | A, B
APC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1:00 | NA | Previous measured value +/- 1% | A, B
APC 3_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
APC 4 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
APC 5_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
APC 6 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
APC 7 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to all Q.H. |  |  | 5 min
APC 8 | Dielectric | Current leakage | Pole, QH | Single pole to each other pole | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
APC 9_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to instrumentation |  |  | 5 min
APC 10 | Discharge test | Discharge | Pole | Single poles | INV | 30 V/turn |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | 60 V/turn
[5 empty] | MQC | 100 V/turn
APC 11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1:00 |  | Previous tests | A, B
APC 12 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
APC 13_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 1000 V | 30 sec. | > 1000 MΩ | C
instrumentation to its poles
APC 14 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
APC 15 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1:00 |  | BC8 | A, B
APC 16 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
Test of collared coil after collaring - AC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6:00 | NA | Previous measured value +/- 1% | A, B
AC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1:00 | NA | Previous measured value +/- 1% | A, B
AC 3_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AC 4 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
AC 5_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
AC 6 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
AC 7 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to all Q.H. |  |  | 5 min
AC 8 | Dielectric | Current leakage | Pole, QH | Single pole to each other pole | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
AC 9_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to instrumentation |  |  | 5 min
AC 10 | Discharge test | Discharge | Pole | Single poles | INV | 30 V/turn |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | 60 V/turn
[5 empty] | MQC | 100 V/turn
AC 11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1:00 |  | Previous tests | A, B
AC 12 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
AC 13_XXXX | insulation resistance | Insulation | BSH | instrumentation to ground | MQC | 1000 V | 30 sec. | > 1000 MΩ | C
instrumentation to its poles
AC 14 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | 450 V (or 40A) |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | 900 V (or 80A)
AC 15 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1:00 |  | BC8 | A, B
AC 16 | Q.H. Dielectric | Current leakage | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | Q.H. to poles |  |  | 5 min
Strip to strip if on different circuits
Test of the magnet after interconnection - AI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AI1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6:00 | NA | Previous measured value +/- 1% | A, B
AI2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1:00 | NA | Previous measured value +/- 1% | A, B
AI3_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AI4 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
AI5_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
AI6 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
AI7 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | VQH[i] | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to all Q.H. |  |  | 5 min
AI8 | Dielectric | Current leakage | Pole, QH | All Poles to ground | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
AI9_XXXX | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | VT[i] | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
AI10 | Discharge test | Discharge | Pole | All poles in series | INV | 60 V/ turn |  | Previous test | D, G, I, K
[5 empty] | MQC | 100 V/ turn
AI11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1:00 |  | Previous tests | A, B
AI12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
AI13_XXXX | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 1 kV | 30 sec. | > 1000 MΩ | C
instrumentation to poles
AI14 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | 900 V (or 80A) |  | Compare initial I and tau with reference | D, H, I
AI15 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1:00 |  | BC8 | A, B
AI16 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | VQH[i] | 30 sec. | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Final Test of the magnet before delivery - FINAL
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
FINAL 1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6:00 | NA | Previous measured value +/- 1% | A, B
FINAL 2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1:00 | NA | Previous measured value +/- 1% | A, B
FINAL 3_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
FINAL 4 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
FINAL 5_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
FINAL 6 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
FINAL 7 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | VtQHaf | 30 sec. | >1000 MΩ; < 10 μA | C
|  |  |  | All poles to all Q.H. |  |  | 5 min
FINAL 8 | Dielectric | Current leakage | Pole, QH | All poles to ground | MQC | VTaf | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
FINAL 9_XXXX | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | VTaf | 30 sec. | >1000 MΩ; < 10 μA | C
5 min
FINAL 10 | Discharge test | Discharge | Pole | All poles in series | INV | 60 V/ turn |  | Previous test | D, G, I, K
[5 empty] | MQC | 100 V/ turn
FINAL 11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1:00 |  | Previous tests | A, B
FINAL 12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | VtQHaf | 30 sec. | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
FINAL 13_XXXX | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 1 kV | 30 sec. | > 1000 MΩ | C
instrumentation to poles
FINAL 14 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | 900 V |  | Compare initial I and tau with reference | D, H, I
FINAL 15 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1:00 |  | BC8 | A, B
FINAL 16 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | VtQHaf | 30 sec. | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
TEST CODE : an identification code for the test
TITLE : Test short name
TYPE : Type of measure
OBJECT : component tested
SH = spot heater, BSH = beam simulation heater, QH = quench heater, TT = temperature sensor
SHORT DESCRIPTION: short description of test if necessary
PRIORITY of test :
a. Mandatory Q.C. [MQC]: the test shall be performed
b. Optional [Opt] : the test can be performed
c. Investigative [Inv]: the test shall be performed as preparatory test of a mandatory Q.C.
TEST CONDITION: i.e. voltage. Two values will be given a set value and a formula to compute the voltage according to previous philosophy above
TIME: duration of test if necessary
TOLERANCE: values inside which the test is considered to be ok
EQUIPMENT: equipment used for the test
the
Large
Hadron
Collider
project
the
Large
Hadron
Collider
project

## Vtest
Summary of tests and of voltages to be applied
| type of magnet | test phase code | Coil/Pole/Aperture Insulation and dielectric to ground | Q.H. Insulation and dielectric to pole and to ground together | Insulation/dielectric instrumentation | Pole Discharge | Q.H. discharge
Series of test | Nb-Ti | L | 200 V | NA | >=1 kV | 120V/turn 5kVmax | NA
|  | BYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | AYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | P | 200 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
| Nb3Sn | PBR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PBI | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAIM | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAI | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PHE | NA | From VQH[i] /10 to VQH[i] /3 | >=1 kV | 120V/turn 5kVmax | 900V 80A
| All magnets | BC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | APC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AI | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | FINAL | Vtaf Table I | VtQHaf Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
L : Test of single layer, Nb-Ti, after curing splice version - L
BYM : Test of pole : 2 layers, Nb-Ti, splice version, before Young modulus – BYM
AYM : Test of pole : 2 layers, Nb-Ti, splice version, after Young modulus – AYM
P : Test of pole : 1 layer test, Nb-Ti, no splice version – P
PBR : Test of pole : Nb3Sn, before reaction – PBR
PAR : Test of pole : Nb3Sn, after reaction – PAR
PBI : Test of pole : Nb3Sn, before impregnation – PBI
PAIM : Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
PAI : Test of pole : Nb3Sn, after impregnation still out of the mould – PAI
PHE : Test of pole : Nb3Sn, after impregnation still out of the mould 0.1 bar He – PHE
BC : Test of collared coil before collaring or pre-collaring – BC
APC : Test of collared coil after pre-collaring – APC
AC : Test of collared coil after collaring – AC
AI : Test of the magnet after interconnection – AI
FINAL : Final Test of the magnet before delivery – FINAL
VQH[i]: Test voltage for the High Voltage test of the Quench Heater to be applied to the ith test
VT[i] : Test voltage for the High Voltage test of the windings to be applied to the ith test

## RDC
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
ELECTRICAL TESTS RESULTS
Test of coil, pole or magnet
Test code PAI1
|  |  | Version no: |  | 1
| Project |  | RMC
| modele : |  | RMC#Copper
PAI1 - RDC Resistance measurement of pole, Nb3Sn after impregnation out of the mould
Appareil utilise : |  | KEITHLEY 2000 multimeter
Alimentation stabilisee de courant
courant aux bornes de chaque pole [A] |  |  |  | 6
|  | Test name |  |  | PAI1
|  | date |  |  | 2013-06-28
|  | Operator name |  |  | FOP
|  | Tamb [°C] |  |  | 22.5
|  | Humidité [%] |  |  | 34.4
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mW]
| I+ | U+ | Lead + | inner coil 1 | 784.03 | 130.6717 | Previous measured | +/- 1%
|  | U+ | Coil 1 upper splice | EES1UO | 777.4 | 129.5667 | Previous measured | +/- 1%
|  | U+ |  | EES1UI | 776.68 | 129.4467 | Previous measured | +/- 1%
|  | U+ | Coil 1 upper layer | EE1U1 | 645.77 | 107.6283 | Previous measured | +/- 1%
|  | U+ |  | EE1U2 | 644.15 | 107.3583 | Previous measured | +/- 1%
|  | U+ |  | EE1U3 | 478.96 | 79.8267 | Previous measured | +/- 1%
|  | U+ |  | EE1U4 | 397.89 | 66.315 | Previous measured | +/- 1%
|  | U+ |  | EE1U5 | 396.22 | 66.0367 | Previous measured | +/- 1%
|  | U+ |  | EE1U6 | 394.42 | 65.7367 | Previous measured | +/- 1%
|  | U+ |  | EE1U7 | 392.77 | 65.4617 | Previous measured | +/- 1%
|  | U+ | Coil 1 lower layer | EE1L1 | 390.93 | 65.155 | Previous measured | +/- 1%
|  | U+ |  | EE1L2 | 389.3 | 64.8833 | Previous measured | +/- 1%
|  | U+ |  | EE1L3 | 387.3 | 64.55 | Previous measured | +/- 1%
|  | U+ |  | EE1L4 | 385.71 | 64.285 | Previous measured | +/- 1%
|  | U+ |  | EE1L5 | 354.56 | 59.0933 | Previous measured | +/- 1%
|  | U+ |  | EE1L6 | 139.47 | 23.245 | Previous measured | +/- 1%
|  | U+ |  | EE1L7 | 137.84 | 22.9733 | Previous measured | +/- 1%
|  | U+ | Coil 1 lower splice | EES1LI | 7.15 | 1.1917 | Previous measured | +/- 1%
|  | U+ |  | EES1LO | 6.38 | 1.0633 | Previous measured | +/- 1%
| I- | U- | Lead - | outer coil 1 | 0.14 | 0.023333 | Previous measured | +/- 1%

## Inductance
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
ELECTRICAL TESTS RESULTS
Test of coil, pole or magnet
Test code PAI3
|  |  |  | Version no: | 1
| QUADRIPÔLE : |  |  | RMC
| modele : |  |  | RMC#Copper
PAI3 - Pole inductance test, Nb3Sn, after impregnation out of the mould
Appareil utilise : |  | BK PRECISION LCR meter 879B
date |  | 2013-06-28
Operator name |  | FOP
Tamb [°C] |  | 23
Humidité [%] |  | 33.7
PAI3 | after impregnation | Inductance [µH]
|  | Mesure réalisée |  | pole 1
|  | Fréquence | 100 Hz | 1412
|  |  | 1000 HZ | 1284
|  |  | 1 0000 Hz | 1068
Tolerances :
Compare single pole among them and to previous assembly

## Insulation
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
ELECTRICAL TESTS RESULTS
Test of coil, pole or magnet
Test code PAI5
|  |  | Version no: |  | 1
| QUADRIPÔLE : |  | RMC
| modele : |  | RMC#Copper
PAI5 - Insulation measurement, Nb3Sn, after impregnation out of the mould
Appareil utilise : |  | MEGGER BM21
| Test name |  | PAI5
| date |  | 2013-06-28
| Operator name |  | FOP
| Tamb [°C] |  | 23
| Humidité [%] |  | 33.7
Insulation resistance |  |  | U[test] | duration | Mesurée | nominale
|  |  | [V] | [s] | [GW] | [GW]
coil1 ---> 3NCS |  |  | 500 | 30 | brd | >1
coil1 ---> 1NCS |  |  | 500 | 30 | 34.5 kΩ | >1
coil1 ---> CP |  |  | 500 | 30 | 13 kΩ | >1
coil1 ---> 1CS |  |  | 500 | 30 | 39.6 kΩ | >1
coil1 ---> 3CS |  |  | 500 | 30 | brd | >1

## Dielectric
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
ELECTRICAL TESTS RESULTS
Test of coil, pole or magnet
Test code PAI5
|  |  | Version no: |  | 1
| QUADRIPÔLE : |  | RMC
| modele : |  | RMC#Copper
PAI5 - Dielectric measurement, Nb3Sn, after impregnation out of the mould
Appareil utilise : |  | MEGGER BM21
| Test name |  | PAI5
| date |  | 2013-06-28
| Operator name |  | FOP
| Tamb [°C] |  | 23
| Humidité [%] |  | 33.7
current leakage |  |  | U[test] | duration | Mesurée | nominale
|  |  | [V] | [min] | [nA] | [µA]
coil1 ---> 3NCS |  |  | 500 | 5 | brd | < 2
coil1 ---> 1NCS |  |  | 500 | 5 | 1.71 mA | < 2
coil1 ---> CP |  |  | 500 | 5 | 1.85 mA | < 2
coil1 ---> 1CS |  |  | 500 | 5 | 1.68 mA | < 2
coil1 ---> 3CS |  |  | 500 | 5 | brd | < 2

## Discharge
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
ELECTRICAL TESTS RESULTS
Test of coil, pole or magnet
Test code PAI4
|  |  |  | Version no: |  | 1
| QUADRIPÔLE : |  |  | RMC
| modele : |  |  | RMC#Copper
PAI4 - Pole interturn discharge, Nb3Sn, after impregnation out of the mould
Appareil utilise : |  |  | SEITZ Impulse tester
Maximum 120V/turn and 5kV/nturn
Graph
Test name |  | PAI4 |  |  | upper layer : xx turns
lower layer : xx turns
Date |  |  |  |  | xx turns/coil |  | measure at x.xxkV, x.xx kV, x.xxkV and xkV
Operator name |  | FO Pincot
Tamb [°C]
Humidité [%]
Time onstant Ƭ |  | Ƭ Mesurée
[µs]
coil1 |  |  |  |  | test OK at 0.x, 0.x et xkv
Graphique