## WBS
Specific Electrical Test result
Test of coil, pole or magnet
Project |  |  | : | MQXC
Model |  |  |  | Aperture_2
[5 empty] | Réalisé par : | Francois-Olivier PINCOT | Date: | 04.05.12
| EDMS no.: |  |  |  | Vérifié par : |  | Version no: | 1.05
LHC Project doc: |  |  |  |  | Approuvé par : |  | Date of last modification: | 12.07.12
the
Large
Hadron
Collider
project

## guideline
Procedure general tests electriques
|  |  |  | Prototype | Version no: |  | 1.05
[5 empty] | Date of last modification: |  | 12.07.12
|  |  |  | Project | MQXC |  | Specific Electrical Test result
Test of coil, pole or magnet
|  |  |  | LHC project doc | Page: | 1 | sur | 3
Equipement tests electriques
| A | KEITHLEY 2000 multimeter
| B | Alimentation stabilisee de courant
| C | MEGGER BM21 / S1-1052
| D | Tektronix TDS 3014
| E | LCR Meter 879B BK Precision
| F | HP 4263B LCR meter
| G | SEITZ Impulse tester
| H | EFACEC Heater discharge Power Supply
| I | Labview data transfert software
| J | Press of Young modulus
| K | Tektronix probe P6015A
| L | Connection cable and/or specific rack
Test of single layer, Nb-Ti, after curing splice version - L
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
Lop1 | Coil to copper wedge | Insulation | Coil | Insulation coil to copper wedge | Opt | 500 V | 30 sec |  | C
L1 | Coil to ground | Insulation | Coil | Insulation of the coil in its curing mould to ground | MQC | 500 V | 30 sec | > 1000 MΩ | C
L2 | RDC | Resistance | Coil | Resistance measurement of coil | MQC | 6A | NA | Ref coil +/- 3% | A, B
L2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
L3 | Inductance | Inductance | Coil | Coil inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
L4 | Coil inter-turn discharge | Discharge | Coil | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
Lop2 | Discharge under stress SP | Discharge | Coil | Discharge while putting the straight part in an insulation mould at P=80MPa | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Lop3 | Discharge under stress ends | Discharge | Coil | Discharge while putting the ends in an insulation mould at P=80MPa decreasing to 0 MPa over the ehads | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Test of pole : 2 layers, Nb-Ti, splice version, before Young modulus - BYM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
BYM5 | Coil inter-turn discharge | Discharge | Coil | Discharge test to identify turn to turn problems | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
BYM12 | Q.H. discharge | Discharge | QH | To be applied if Q.H are under coils and already assembled | Opt | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
Test of pole : 2 layers, Nb-Ti, splice version, after Young modulus - AYM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AYM1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Inner layer +outer layer +/- 2% | A, B
AYM2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
AYM2_XXXX | RDC instrumentation | Resistance |  | RDC of each component | MQC | 1A / 4 wires | NA | Compare to previous component measurement +/- 2% | A, B
AYM3 | Inductance | Inductance | Pole | Pole inductance measurement | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref pole] +/- 2% | E
AYM4 | Insulation resistance | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | MQC | 1 kV | 30 sec | > 1000 MΩ | C
AYM5 | Dielectric | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | MQC | 2.4 kV |  | < 5 μA
AYM6 | Coil discharge test | Discharge | Pole | Discharge test to identify turn to turn problems | INV | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AYM7 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
AYM8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
AYM9 | Insulation resistance | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | MQC | 1 kV | 30 sec | < 1000 MΩ | C
AYM10 | Dielectric | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | MQC | 2.4 kV |  | < 5 μA
Test of pole : 1 layer test, Nb-Ti, no splice version - P
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
Pop1 | Pole to copper wedge | Insulation | Pole | Insulation pole to copper wedge | Opt | 500 V | 30 sec |  | C
P1 | Pole to ground | Insulation | Pole | Insulation of the pole in its curing mould to ground | MQC | 500 V | 30 sec | > 1000 MΩ | C
P2 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
P3 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled
P3_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
P4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
P5 | Coil inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
Pop2 | Discharge under stress SP | Discharge | Pole | Discharge while putting the straight part in an insulation mould at P=80MPa | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
Pop3 | Discharge under stress ends | Discharge | Pole | Discharge while putting the ends in an insulation mould at P=80MPa decreasing to 0 MPa over the ehads | Opt | Max [120 V/turn, 5kV/nturn] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
P6 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled | MQC | 1 kV | 30 sec | < 1000 MΩ | C
P7 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
P8 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
P9 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled | MQC | 1 kV | 30 sec | < 1000 MΩ | C
P10 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
Test of pole : Nb3Sn, before reaction - PBR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
PBR2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBR3 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PBR4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBR5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before reaction | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PBR6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before reaction | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PBR7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before reaction | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PBR8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before reaction
Test of pole : Nb3Sn, after reaction - PAR
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAR1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAR2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled after reaction
PAR2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAR3 | Pole to ground | Insulation | Pole | Insulation of the pole in its reaction mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAR4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAR5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled after reaction | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAR6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled after reaction | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAR7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled after reaction | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAR8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled after reaction
Test of pole : Nb3Sn, before impregnation - PBI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PBI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PBI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PBI2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PBI3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PBI4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PBI5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PBI6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PBI7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PBI8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAIM1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAIM2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAIM2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAIM3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAIM4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAIM5 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAIM6 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAIM7 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAIM8 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
Test of pole : Nb3Sn, after impregnation still out of the mould - PAI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PAI1 | RDC | Resistance | Pole | Resistance measurement of pole | MQC | 6A | NA | Ref coil +/- 3% | A, B
PAI2 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAI2_XXXX | RDC | Resistance |  | Resistance measurement of intrumentation | MQC | 1A / 4 wires | NA | Reference value for future measurements | A, B
PAI3 | Pole to ground | Insulation | Pole | Insulation of the pole in its impregnation mould to ground | Inv | 100 V | 30 sec | > 1000 MΩ | C
PAI4 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PAI5 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | Inv | Max [(120 V/turn, 5kV/nturn)/4] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | Inv | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
PAI6 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | Inv | 1 kV | 30 sec | < 1000 MΩ | C
PAI7 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
PAI8 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled before impregnation | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
PAI9 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 1 kV | 30 sec | < 1000 MΩ | C
PAI10 | RDC Q.H. | Resistance | QH | Measurement of Q.H. RDC after assembly | MQC | 1A | NA | Compare to previous component measurement +/- 2% | A, B
To be applied if Q.H are under coils and already assembled before impregnation
PAI11 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 1 kV | 5 min | < 2 μA | C
[5 empty] | INV | 3 kV |  | < 5 μA
[5 empty] | MQC | 6.6 kV |  | < 10 μA
Test of pole : Nb3Sn, after impregnation still out of the mould 0.1 bar He - PHE
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
PHE1 | Inductance | Inductance | Pole | Pole inductance test | MQC | 100 Hz, 1 KHz, 10 KHz | NA | L[ref coil] +/- 2% | E
PHE2 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | 20 V/turn | 10 puls | Ref coil τ +/- 1% | D, G, I, K
[5 empty] | MQC | 40 V/turn
[5 empty] | MQC | 60 V/turn
PHE3 | Pole inter-turn discharge | Discharge | Pole | Discharge test to identify turn to turn problems | MQC | Max [(120 V/turn, 5kV/nturn)] | 10 puls | Ref coil τ +/- 1% | D, G, I, K
PHE4 | Insulation resistance Q.H. to coil | Insulation | Pole, QH | To be applied if Q.H are between coils and already assembled before impregnation | MQC | 250 V | 30 sec | < 1000 MΩ | C
PHE5 | Dielectric Q.H. to coil pre-test | Current leakage | Pole, QH | To be applied if Q.H are between coils and already assembled QH-> coil | INV | 200 V | 5 min | < 1 μA | C
[5 empty] | INV | 500 V |  | < 2.5 μA
[5 empty] | MQC | 1.7 kV |  | < 10 μA
Test of collared coil before collaring or pre-collaring - BC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
BC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
BC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
BC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
BC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
BC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
BC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
BC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.4 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
BC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.4 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
BC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.4 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
BC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.4 kV | 5 min | < 10 μA | C
All poles to instrumentation
BC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
BC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
BC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.4 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
BC9_XXXX | insulation resistance | Insulation | BSH | instrumentation to ground | MQC | 3.4 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
BC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
BC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
BC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.4 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of collared coil after pre-collaring - APC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
APC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
APC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
APC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
APC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
APC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
APC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
APC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.2 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
APC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.2 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
APC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.2 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
APC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.2 kV | 5 min | < 10 μA | C
All poles to instrumentation
APC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
APC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
APC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.2 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
APC9_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.2 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
APC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
APC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
APC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.2 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of collared coil after collaring - AC
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AC1 | RDC | Resistance | Pole | RDC of each poles and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
AC2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
AC2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AC3 | Capacitance | Capacitance | Pole, QH | Q.H. to its pole on which they are assembled | MQC | NA | NA | Compare single poles among them and to previous assembly | E
Single pole to ground
Single pole to single pole
All poles to ground
AC3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare single result among them and to previous assembly | E
AC4 | Inductance | Inductance | Pole | L of each poles | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare single poles among them and to previous assembly | E
L of all poles temporarily in series
AC5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
Single pole to each other pole
AC5_XXXX | Insulation resistance | Insulation |  | All Poles + instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
AC6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
Single pole to each other pole
AC6_XXXX | Dielectric | Current leakage |  | All Poles + instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
AC7 | Discharge test | Discharge | Pole | Single poles | INV | Max [(120 V/turn, 5kV/nturn)/4] |  | Previous test | D, G, I, K
|  |  |  | all poles temporarily in series | INV | Max [(120 V/turn, 5kV/nturn)/2]
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AC8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
AC9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
AC9_XXXX | insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to its poles
AC10 | Q.H. discharge | Discharge | QH | To be applied if Q.H are between coils and already assembled | INV | Max [(850V, 80A)/2] |  | Compare initial I and tau with reference | D, H, I
[5 empty] | MQC | Max [850V, 80A]
AC11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
AC12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Test of the magnet after interconnection - AI
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
AI1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
AI2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
AI2_XXXX | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
AI3 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
AI3_XXXX | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
AI4 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
AI5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
All poles to ground
AI5_XXXX | Insulation resistance | Insulation |  | All Poles + All instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
AI6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
All poles to ground
AI6_XXXX | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
AI7 | Discharge test | Discharge | Pole | All poles in series | INV | Max [(120 V/turn, 5kV/nturn)/2] |  | Previous test | D, G, I, K
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
AI8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
AI9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
AI9_XXXX | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to poles
AI10 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
AI11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
AI12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
Final Test of the magnet before delivery - FINAL
Test code | Title | Type | Object | Short description | Priority | Test condition | Time [sec] | Tolerances | Equipment
FINAL1 | RDC | Resistance | Pole | RDC of poles in serie and Vtaps | MQC | 6 A | NA | Previous measured value +/- 1% | A, B
FINAL2 | RDC QH | Resistance | QH | RDC of each Q.H. | MQC | 1 A | NA | Previous measured value +/- 1% | A, B
FINAL2_MQXC | RDC instrumentation | Resistance |  | RDC of instrumentation | MQC | 1A / 4 wires | NA | Previous measured value +/- 1% | A, B
FINAL3 | Capacitance | Capacitance | Pole, QH | Q.H. to poles | MQC | NA | NA | Compare to previous assembly | E
All poles to ground
FINAL3_MQXC | Capacitance | Capacitance |  | Instrumentation to poles between which they are | MQC | NA | NA | Compare to previous assembly | E
FINAL4 | Inductance | Inductance | Pole | L of all poles in series | MQC | 100 Hz, 1 KHz, 10 KHz |  | Compare to previous assembly | E
FINAL5 | Insulation resistance | Insulation | Pole, QH | Poles + Q.H to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to all Q.H.
All poles to ground
FINAL5_MQXC | Insulation resistance | Insulation |  | All Poles + All instrumentation to ground | MQC | 1.0 kV | 30 sec | > 1 MΩ | C
All poles to instrumentation
FINAL6 | Dielectric | Current leakage | Pole, QH | Poles + Q.H to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to all Q.H.
All poles to ground
FINAL6_MQXC | Dielectric | Current leakage |  | All Poles + All instrumentation to ground | MQC | 6.0 kV | 5 min | < 10 μA | C
All poles to instrumentation
FINAL7 | Discharge test | Discharge | Pole | All poles in series | INV | Max [(120 V/turn, 5kV/nturn)/2] |  | Previous test | D, G, I, K
[5 empty] | MQC | Max [120 V/turn, 5kV/nturn]
FINAL8 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | Opt | 1 A |  | Previous tests | A, B
FINAL9 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
Q.H. to poles
Strip to strip if on different circuits
FINAL9_MQXC | Insulation resistance | Insulation |  | instrumentation to ground | MQC | 3.0 kV | 1 min | > 1000 MΩ | C
instrumentation to poles
FINAL10 | Q.H. discharge | Discharge | QH | Discharge each single Q.H. | MQC | Max [850V, 80A] |  | Compare initial I and tau with reference | D, H, I
FINAL11 | Q.H. RDC | Resistance | QH | Resistance each Q.H. | MQC | 1 A |  | BC8 | A, B
FINAL12 | Q.H. insulation resistance | Insulation | QH | Q.H. to ground | MQC | 3.0 kV | 1 min | BC9 | C
Q.H. to poles
Strip to strip if on different circuits
1) TEST CODE : an identification code for the test
2) TITLE : Test short name
3) TYPE : Type of measure
4) OBJECT : component tested
SH = spot heater, BSH = beam simulation heater, QH = quench heater, TT = temperature sensor
5) SHORT DESCRIPTION: short description of test if necessary
6) PRIORITY of test :
a. Mandatory Q.C. [MQC]: the test shall be performed
b. Optional [Opt] : the test can be performed
c. Investigative [Inv]: the test shall be performed as preparatory test of a mandatory Q.C.
7) TEST CONDITION: i.e. voltage. Two values will be given a set value and a formula to compute the voltage according to previous philosophy above
8) TIME: duration of test if necessary
9) TOLERANCE: values inside which the test is considered to be ok
10) EQUIPMENT: equipment used for the test
the
Large
Hadron
Collider
project
the
Large
Hadron
Collider
project

## Vtest
Summary of tests and of voltages to be applied
| type of magnet | test phase code | Coil/Pole/Aperture Insulation and dielectric to ground | Q.H. Insulation and dielectric to pole and to ground together | Insulation/dielectric instrumentation | Pole Discharge | Q.H. discharge
Series of test | Nb-Ti | L | 200 V | NA | >=1 kV | 120V/turn 5kVmax | NA
|  | BYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | AYM | NA | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | P | 200 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
| Nb3Sn | PBR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAR | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PBI | 50 V | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAIM | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PAI | VT[i] Table I | VQH[i] Table II | >=1 kV | 120V/turn 5kVmax | 900V 80A
|  | PHE | NA | From VQH[i] /10 to VQH[i] /3 | >=1 kV | 120V/turn 5kVmax | 900V 80A
| All magnets | BC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | APC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AC | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | AI | VT[i] Table I | VQH[i] Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
|  | FINAL | Vtaf Table I | VtQHaf Table II | >=1 kV | 100V/turn 4.25kVmax | 900V 80A
L : Test of single layer, Nb-Ti, after curing splice version - L
BYM : Test of pole : 2 layers, Nb-Ti, splice version, before Young modulus – BYM
AYM : Test of pole : 2 layers, Nb-Ti, splice version, after Young modulus – AYM
P : Test of pole : 1 layer test, Nb-Ti, no splice version – P
PBR : Test of pole : Nb3Sn, before reaction – PBR
PAR : Test of pole : Nb3Sn, after reaction – PAR
PBI : Test of pole : Nb3Sn, before impregnation – PBI
PAIM : Test of pole : Nb3Sn, after impregnation still in the mould - PAIM
PAI : Test of pole : Nb3Sn, after impregnation still out of the mould – PAI
PHE : Test of pole : Nb3Sn, after impregnation still out of the mould 0.1 bar He – PHE
BC : Test of collared coil before collaring or pre-collaring – BC
APC : Test of collared coil after pre-collaring – APC
AC : Test of collared coil after collaring – AC
AI : Test of the magnet after interconnection – AI
FINAL : Final Test of the magnet before delivery – FINAL
VQH[i]: Test voltage for the High Voltage test of the Quench Heater to be applied to the ith test
VT[i] : Test voltage for the High Voltage test of the windings to be applied to the ith test

## RDC
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L2 / AYM1 / BC1 / APC1 / AC1 / AI1 / FINAL1
|  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
L2 - RDC Test of single layer, Nb-Ti, after curing splice version
AYM1 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young
BC1 - RDC Mesure de resistance de chaque pole et des V-taps avant pre-collaring
APC1 - RDC Mesure de resistance de chaque pole et des V-taps apres pre-collaring
AC1 - RDC Mesure de resistance de chaque pole et des V-taps apres collaring
AI1 - RDC Mesure de resistance de chaque pole et des V-taps apres interconnection
FINAL1 - RDC Mesure de resistance de chaque pole et des V-taps avant livraison
Appareil utilise : |  | KEITHLEY 2000 multimeter
Alimentation stabilisee de courant
courant aux bornes de chaque pole [A] |  |  |  | 6
|  | Nom du test |  |  | FINAL1 |  | Vtaps protec
|  | Tests effectués le : |  |  | 2012-12-21 |  | 2013-01-15
|  | Nom de l'opérateur : |  |  | FOP |  | FOP
|  | Tamb [°C] |  |  | 20.3 |  | 20.7
|  | Humidité [%] |  |  | 29.5 |  | 19.1
All coils in serie
R[dc] Dipole Voltage Taps |  |  | Layer | nom | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [mV] | [mW] | [mV] | [mW] | [mW]
| I+ | U+ | inner coil9 | inner coil 9 | 3816 | 636 | 3771 | 628.5 | Previous measured | +/- 1%
|  | U+ |  | EEHI91 | 3409 | 568.1667 | 3373 | 562.1667 | Previous measured | +/- 1%
|  | U+ |  | EEHI92 | 3409 | 568.1667 | 3373 | 562.1667 | Previous measured | +/- 1%
|  | U+ |  | EEJI91 | 3402 | 567 | 3365 | 560.8333 | Previous measured | +/- 1%
|  | U+ |  | EEJI92 | 3402 | 567 | 3365 | 560.8333 | Previous measured | +/- 1%
|  | U+ |  | EEI91 | 3402 | 567 | 3365 | 560.8333 | Previous measured | +/- 1%
|  | U+ |  | EEI92 | 3402 | 567 | 3365 | 560.8333 | Previous measured | +/- 1%
|  | U+ | outer coil9 | EEO91 | 3401 | 566.8333 | 3364 | 560.6667 | Previous measured | +/- 1%
|  | U+ |  | EEO92 | 3401 | 566.8333 | 3364 | 560.6667 | Previous measured | +/- 1%
|  | U+ | splice 9-3 | EESO91 | 2858 | 476.3333 | 2822 | 470.3333 | Previous measured | +/- 1%
|  | U+ |  | EESO92 | 2858 | 476.3333 | 2822 | 470.3333 | Previous measured | +/- 1%
|  | U+ |  | EESI31 | 2858 | 476.3333 | 2822 | 470.3333 | Previous measured | +/- 1%
|  | U+ |  | EESI32 | 2858 | 476.3333 | 2822 | 470.3333 | Previous measured | +/- 1%
|  | U+ | inner coil3 | EEHI31 | 2470 | 411.6667 | 2435 | 405.8333 | Previous measured | +/- 1%
|  | U+ |  | EEHI32 | 2470 | 411.6667 | 2435 | 405.8333 | Previous measured | +/- 1%
|  | U+ |  | EEJI31 | 2461 | 410.1667 | 2427 | 404.5 | Previous measured | +/- 1%
|  | U+ |  | EEJI32 | 2461 | 410.1667 | 2427 | 404.5 | Previous measured | +/- 1%
|  | U+ |  | EEI31 | 2460 | 410 | 2426 | 404.3333 | Previous measured | +/- 1%
|  | U+ |  | EEI32 | 2460 | 410 | 2426 | 404.3333 | Previous measured | +/- 1%
|  | U+ | outer coil3 | EEO31 | 2460 | 410 | 2426 | 404.3333 | Previous measured | +/- 1%
|  | U+ |  | EEO32 | 2460 | 410 | 2426 | 404.3333 | Previous measured | +/- 1%
|  | U+ | splice 3-7 | EESO31 | HS | #VALUE! | 1885 | 314.1667 | Previous measured | +/- 1%
|  | U+ |  | EESO32 | 1918 | 319.6667 | 1885 | 314.1667 | Previous measured | +/- 1%
|  | U+ |  | EESO71 | 1918 | 319.6667 | HS | #VALUE! | Previous measured | +/- 1%
|  | U+ |  | EESO72 | 1918 | 319.6667 | 1885 | 314.1667 | Previous measured | +/- 1%
|  | U+ | outer coil7 | EEO71 | 1373 | 228.8333 | 1341 | 223.5 | Previous measured | +/- 1%
|  | U+ |  | EEO72 | 1372 | 228.6667 | 1341 | 223.5 | Previous measured | +/- 1%
|  | U+ | inner coil7 | EEI71 | 1372 | 228.6667 | 1341 | 223.5 | Previous measured | +/- 1%
|  | U+ |  | EEI72 | 1372 | 228.6667 | 1341 | 223.5 | Previous measured | +/- 1%
|  | U+ | splice 7-8 | EESI71 | 974.9 | 162.4833 | 945 | 157.5 | Previous measured | +/- 1%
|  | U+ |  | EESI72 | 974.9 | 162.4833 | 945.8 | 157.6333 | Previous measured | +/- 1%
|  | U+ |  | Lead in7out8 | 974.9 | 162.4833 | 945.7 | 157.6167 | Previous measured | +/- 1%
|  | U+ |  | EESO81 | 974.9 | 162.4833 | 944.7 | 157.45 | Previous measured | +/- 1%
|  | U+ |  | EESO82 | 974.9 | 162.4833 | 944.8 | 157.4667 | Previous measured | +/- 1%
|  | U+ | outer coil8 | EEO81 | 443.6 | 73.9333 | 403 | 67.1667 | Previous measured | +/- 1%
|  | U+ |  | EEO82 | 443.6 | 73.9333 | 403 | 67.1667 | Previous measured | +/- 1%
|  | U+ | inner coil8 | EEI81 | 432.7 | 72.1167 | 402.6 | 67.1 | Previous measured | +/- 1%
|  | U+ |  | EEI82 | 432.7 | 72.1167 | 402.6 | 67.1 | Previous measured | +/- 1%
|  | U+ |  | inner coil 8 | 48.5 | 8.0833 | 0.2 | 0.033333 | Previous measured | +/- 1%

## heater RDC
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES SPECIFIQUES
Test of coil, pole or magnet
Test code AYM2-8 / BC2-8-11 / AC2-8-11 / APC2-8-11 / AI2-8-11 / FINAL2-8-11
|  |  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
AYM2 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young
AYM8 - RDC Mesure de resistance de chaque pole et des V-taps apres mesure module d'Young et apres test decharge
BC 2 - RDC - Mesure initiale de resistance de chaque Quench Heater (QH) - avant pre-collaring
BC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - avant pre-collaring
BC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge
APC 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres pre-collaring
APC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres pre-collaring
APC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres pre-collaring
AC 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres collaring
AC 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres collaring
AC 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres collaring
AI 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - apres interconnection
AI 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - apres interconnection
AI 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - apres interconnection
FINAL 2 - RDC - Mesure de resistance initiale de chaque Quench Heater (QH) - avant livraison
FINAL 8 - RDC - Mesure de resistance de chaque Quench Heater (QH) avant test de decharge - avant livraison
FINAL 11 - RDC - Mesure de resistance de chaque Quench Heater (QH) apres test de decharge - avant livraison
Appareil utilise : |  |  | KEITHLEY 2000 multimeter
Alimentation stabilisee de courant
courant aux bornes de chaque QH [A] |  |  |  | 1
|  | Nom du test |  |  | FINAL2 - initial |  | FINAL8 - bf discharge |  | FINAL11 - af discharge
|  | Tests effectués le : |  |  | 2012-12-21 |  |  |  | 2013-01-07
|  | Nom de l'opérateur : |  |  | FOP |  |  |  | FOP
|  | Tamb [°C] |  |  | 21.5 |  |  |  | 21.2
|  | Humidité [%] |  |  | 28.3 |  |  |  | 26.3
R[dc] Dipole Voltage Taps |  |  |  | nom | Mesurée |  | Mesurée |  | Mesurée |  | nominale
[5 empty] | Tension | Resistance | Tension | Resistance | Tension | Resistance | Resistance
[5 empty] | [V] | [W] | [V] | [W] | [V] | [W] | [W]
Quench Heater
| I+ |  | U+ | YTI7L | 7.402 | 7.402 |  | 0 | 7.391 | 7.391 | Previous measured | +/- 1%
| I- |  | U- | YTO7L | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI7R | 7.364 | 7.364 |  | 0 | 7.352 | 7.352 | Previous measured | +/- 1%
| I- |  | U- | YTO7R | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI3L | 7.381 | 7.381 |  | 0 | 7.37 | 7.37 | Previous measured | +/- 1%
| I- |  | U- | YTO3L | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI3R | 7.39 | 7.39 |  | 0 | 7.379 | 7.379 | Previous measured | +/- 1%
| I- |  | U- | YTO3R | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI8L | 7.327 | 7.327 |  | 0 | 7.391 | 7.391 | Previous measured | +/- 1%
| I- |  | U- | YTO8L | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L1 | 7.774 | 7.774 |  | 0 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L1 | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8L2 | 7.771 | 7.771 |  | 0 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8L2 | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI8R | 7.372 | 7.372 |  | 0 | 7.361 | 7.361 | Previous measured | +/- 1%
| I- |  | U- | YTO8R | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R1 | 7.817 | 7.817 |  | 0 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R1 | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | EYTI8R2 | 7.817 | 7.817 |  | 0 |  | 0 | Previous measured | +/- 1%
| I- |  | U- | EYTO8R2 | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI9L | HS | #VALUE! |  | 0 | HS | #VALUE! | Previous measured | +/- 1%
| I- |  | U- | YTO9L | [6 empty] | Previous measured | +/- 1%
| I+ |  | U+ | YTI9R | 7.369 | 7.369 |  | 0 | 7.358 | 7.358 | Previous measured | +/- 1%
| I- |  | U- | YTO9R | [6 empty] | Previous measured | +/- 1%
Beam Simulation Heater
| I+ |  | U+ | BSI73 | 1.724 | 1.724 | 1.722 | 1.722
| I- |  | U- | BSO73
| I+ |  | U+ | EBSI73 | 1.971 | 1.971 | 1.968 | 1.968
| I- |  | U- | EBSO73
| I+ |  | U+ | BSI73 | 0.8743 | 0.8743 | 0.906 | 0.906
| I- |  | U- | BSM73
| I+ |  | U+ | EBSI73 | 1.1523 | 1.1523 | 1.426 | 1.426
| I- |  | U- | EBSM73
| I+ |  | U+ | BSI38 | 1.738 | 1.738
| I- |  | U- | BSO38
| I+ |  | U+ | EBSI38 | 2.011 | 2.011
| I- |  | U- | EBSO38
| I+ |  | U+ | BSI89 | 1.778 | 1.778
| I- |  | U- | BSO89
| I+ |  | U+ | EBSI89 | 2.082 | 2.082
| I- |  | U- | EBSO89
| I+ |  | U+ | BSI97 | 1.732 | 1.732
| I- |  | U- | BSO97
| I+ |  | U+ | EBSI97 | 2.008 | 2.008
| I- |  | U- | EBSO97
Spot Heater
| I+ |  | U+ | SHI3 | 6.935 | 17.3375 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SHO3 |  |  | Previous measured | +/- 1%
| I+ |  | U+ | SHI9 | 7.095 | 17.7375 | Previous measured | +/- 1% | 0.4 | A
| I- |  | U- | SHO9 |  |  | Previous measured | +/- 1%
Temperature sensor (4 wires measure)
|  |  |  | TTI3 |  | HS
|  |  |  | TTI9 |  | HS
|  |  |  | TTS39 |  | 912.15
|  |  |  | TTSI3 |  | 998.32
|  |  |  | TTCI |  | 944.82
|  |  |  | TTCO |  | 940.3
|  |  |  | TTL |  | 983.95

## Capacitance
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code BC3 / APC3 / AC3 / AI3 / FINAL3
|  |  | Version no: | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
BC3 - Capacitance - Mesure de la capacite des poles avant pre-collaring
APC3 - Capacitance - Mesure de la capacite des poles apres pre-collaring
AC3 - Capacitance - Mesure de la capacite des poles apres collaring
AI3 - Capacitance - Mesure de la capacite des poles apres interconnection
FINAL3 - Capacitance - Mesure de la capacite des poles avant livraison
Appareil utilise : |  | LCR Meter 879B BK Precision
Mesures realisees a une frequence de (Hz): |  |  | 1000
| Nom du test |  | FINAL3 - Final
| Tests effectués le : |  | 2012-12-21
| Nom de l'opérateur : |  | FOP
| Tamb [°C] |  | 22.1
| Humidité [%] |  | 29.4
Capacitance |  |  | Mesurée
[nF]
QH7L ---> all poles |  |  | 3.91
QH7R ---> all poles |  |  | 3.95
QH3L ---> all poles |  |  | 3.98
QH3R ---> all poles |  |  | 3.93
QH8L ---> all poles |  |  | 4.18
QH8R ---> all poles |  |  | 4.13
QH9L ---> all poles |  |  | 4.03
QH9R ---> all poles |  |  | 3.98
All poles ---> Ground |  |  | 17.47
Tolerances :
Compare single pole among them and to previous assembly

## Inductance
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L3 / AYM3 / BC4 / APC4 / AC4 / AI4 / FINAL4
|  |  | Version no: | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
L3 -Test of single layer, Nb-Ti, after curing splice version
AYM3 - Mesure de l'inductance des poles apres mesure module d'Young
BC4 - Mesure de l'inductance des poles avant pre-collaring
APC4 - Mesure de l'inductance des poles apres pre-collaring
AC4 - Mesure de l'inductance des poles apres collaring
AI4 - Mesure de l'inductance des poles apres interconnection
FINAL4 - Mesure de l'inductance des poles avant livraison
Appareil utilise : |  | BK PRECISION LCR meter 879B
Tests effectués le : |  | 2012-12-21
Nom de l'opérateur : |  | FOP
Tamb [°C] |  | 22.7
Humidité [%] |  | 27.6
Inductance [mH]
Test Code |  | FINAL
Mesure réalisée |  | all poles in serie
Fréquence | 100 Hz | 8.26
| 1000 HZ | 7.1
| 1 0000 Hz | 4.31
Tolerances :
Compare single pole among them and to previous assembly

## Insulation
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM 4-9 / BC5 / APC5 / AC5 / AI5 / FINAL5
|  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
AYM4 - Mesure de resistance d'isolation des poles apres module d'Young
AYM9 - Mesure de resistance d'isolation des poles apres module d'Young apres test decharge
BC5 - Mesure de resistance d'isolation des poless avant pre-collaring
APC5 -Mesure de resistance d'isolation des poles apres pre-collaring
AC5 - Mesure de resistance d'isolation des poles apres collaring
AI5 - Mesure de resistance d'isolation des poles apres interconnection
FINAL5 - Mesure de resistance d'isolation des poles avant livraison
Appareil utilise : |  | MEGGER BM21
| Nom du test |  | FINAL5 - Final |  |  | FINAL5 - Vtaps protec
| Tests effectués le : |  | 2012-12-21 |  |  | 2013-01-15
| Nom de l'opérateur : |  | FOP |  |  | FOP
| Tamb [°C] |  | 21.8 |  |  | 20.7
| Humidité [%] |  | 29.6 |  |  | 19.1
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [s] | [GW] | [V] | [s] | [GW] | [GW]
All poles + all QHs ---> Ground |  |  | 1000 | 30 | 47.7 | 1000 | 30 |  | >1
all poles ---> all QHs |  |  | 1000 | 30 | 76.5 | 1000 | 30 |  | >1
all poles ---> Ground |  |  | 1000 | 30 | 44.5 | 1000 | 30 | 41.7 | >1

## Dielectric
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM5-10 / BC6 / APC6 / AC6 / AI6 / FINAL6
|  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
AYM5 - Mesure du courant de fuite dans le dielectrique apres module d'Young
AYM10 - Mesure du courant de fuite dans le dielectriques apres module d'Young et apres test decharge
BC6 - Mesure du courant de fuite dans le dielectrique avant pre-collaring
APC6 - Mesure du courant de fuite dans le dielectriques apres pre-collaring
AC6 - Mesure du courant de fuite dans le dielectrique apres collaring
AI6 - Mesure du courant de fuite dans le dielectrique apres interconnection
FINAL6 - Mesure du courant de fuite dans le dielectrique avant livraison
Appareil utilise : |  | MEGGER BM21
| Nom du test |  | FINAL6 - Final |  |  | FINAL6 - Vtaps protec
| Tests effectués le : |  | 2012-12-21 |  |  | 2013-01-15
| Nom de l'opérateur : |  | FOP |  |  | FOP
| Tamb [°C] |  | 21.8 |  |  | 20.7
| Humidité [%] |  | 29.6 |  |  | 19.1
Insulation resistance |  |  | U[test] | duration | Mesurée | U[test] | duration | Mesurée | nominale
|  |  | [V] | [min] | [µA] | [V] | [min] | [µA] | [µA]
All poles + all QHs ---> Ground |  |  | 2400 | 5 | 0.0149 | 2400 | 5 |  | < 10
all poles ---> all QHs |  |  | 2400 | 5 | 0.007 | 2400 | 5 |  | < 10
All poles ---> Ground |  |  | 2400 | 5 | 0.0142 | 2400 | 5 | 0.0187 | < 10

## Discharge
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code L4 / AYM6 / BC7 / APC7 / AC7 / AI7 / FINAL7
|  |  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
L4 - Test of single layer, Nb-Ti, after curing splice version
AYM6 - Discharge test - Mesure du temps de decharge des poles apres module d'Young
BC7 - Discharge test - Mesure du temps de decharge des poles avant pre-collaring
APC7 - Discharge test - Mesure du temps de decharge des poles apres pre-collaring
AC7 - Discharge test - Mesure du temps de decharge des poles apres collaring
AI7 - Discharge test - Mesure du temps de decharge des poles apres interconnection
FINAL7 - Discharge test - Mesure du temps de decharge des poles avant livraison
Appareil utilise : |  |  | SEITZ Impulse tester
Maximum 120V/turn and 5kV/nturn
| inner coil : 17 turns |  | measure at 1.7kV and 2.04kV
| outer coil : 19 turns |  | measure at 1.9kV and 2.3kV
Graphique
Nom du test |  | FINAL7 |  |  | 36 turns by pole |  | measure at 3.6kV and 4.3kV
|  | Final |  |  | after interconnection measure at 5.0kV
Tests effectués le : |  | 2013-01-07
Nom de l'opérateur : |  | FOP
Tamb [°C] |  | 20.7
Humidité [%] |  | 26.8
Constante de temps Ƭ |  | Ƭ Mesurée
[µs]
Pole 7 |  | All poles
Pole 3
Pole 8
Pole 9
Graphique

## Heater insulation
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES SPECIFIQUES
Test of coil, pole or magnet
Test code BC9-12 / APC9-12 / AC9-12 / AI9-12 / FINAL 9-12
|  |  | Version no: |  | 1.05
| QUADRIPÔLE : |  | MQXC
| modele : |  | Aperture_2
BC9 - Mesure de resistance d'isolation des heaters avant pre-collaring
BC12 - Mesure de resistance d'isolation des heaters apres test decharge avant pre-collaring
APC9 - Mesure de resistance d'isolation des heaters apres pre-collaring
ACP12 - Mesure de resistance d'isolation des heaters apres test decharge apres pre-collaring
AC9 - Mesure de resistance d'isolation des heaters apres collaring
AC12 - Mesure de resistance d'isolation des heaters apres test decharge apres collaring
AI9 - Mesure de resistance d'isolation des heaters apres interconnection
AI12 - Mesure de resistance d'isolation des heaters apres test decharge apres interconnection
FINAL9 - Mesure de resistance d'isolation des heaters avant livraison
FINAL12 - Mesure de resistance d'isolation des heaters apres test decharge avant livraison
Appareil utilise : |  | MEGGER BM21
| Tests effectués le : |  | 2013-01-07
| Nom de l'opérateur : |  | FOP
| Tamb [°C] |  | 21.2
| Humidité [%] |  | 26.3
| Nom du test |  | FINAL12 - Final
Insulation resistance |  |  | U[test] | duration | Mesurée
|  |  | [V] | [s] | [GW]
All QHs --> Ground |  |  | 3000 | 60 | 126
All BSHs --> Ground |  |  | 3000 | 60 | 102
All SHs --> Ground |  |  | 2500 | 60 | 2340
QH 9L --> all poles |  |  | 3000 | 60 | 833
QH 9R --> all poles |  |  | 3000 | 60 | 986
QH 8L --> all poles |  |  | 3000 | 60 | 703
QH 8R --> all poles |  |  | 3000 | 60 | 788
QH 3L --> all poles |  |  | 3000 | 60 | 722
QH 3R --> all poles |  |  | 3000 | 60 | 737
QH 7L --> all poles |  |  | 3000 | 60 | 828
QH 7R --> all poles |  |  | 3000 | 60 | 999
BSH7-3 -->all poles |  |  | 3000 | 60 | 115
BSH3-8 --> all poles |  |  | 3000 | 60 | 129
BSH8-9 --> all poles |  |  | 3000 | 60 | 132
BSH9-7 --> all poles |  |  | 3000 | 60 | 169
SH3 --> all poles |  |  | 3000 | 60 | 4330
SH9 --> all poles |  |  | 3000 | 60 | 4240
TT3 --> Ground |  |  | 1000 | 60 | HS
TT9 --> Ground |  |  | 1000 | 60 | HS
TT3 --> All poles |  |  | 1000 | 60 | HS
TT9 --> All poles |  |  | 1000 | 60 | HS
TTS39 --> Ground |  |  | 1000 | 60 | 2360
TTSI3 --> Ground |  |  | 1000 | 60 | 3000
TTS39 --> All poles |  |  | 1000 | 60 | 2040
TTSI3 --> All poles |  |  | 1000 | 60 | 3000
Poles --> Ground - TTCI/CO/L connect gnd |  |  | 1000 | 60 | 78.1
Instru --> Ground - TTCI/CO/L connect gnd |  |  | 1000 | 60 | 80.8

## QH discharge
Magnets, Superconductors and Cryostats (MSC)
Magnet Design and Technology (MDT)
BATIMENT 927
Report resultats TESTS ELECTRIQUES
Test of coil, pole or magnet
Test code AYM7 / BC10 / APC10 / AC10 / AI10 / FINAL10
|  |  |  | Version no: | 1.05
| QUADRIPÔLE : |  |  | MQXC
| modele : |  |  | Aperture_2
AYM7 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres module d'Young
BC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH avant pre-collaring
APC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres pre-collaring
AC10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres collaring
AI10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH apres interconnection
FINAL10 - QH Discharge test - Mesure constante de temps lors de la decharge des QH avant livraison
Appareil utilise : |  |  | EFACEC Heater discharge Power Supply
Tektronix TDS 3014
Labview data transfert software
condition test [V] |  |  |  | 450
heater discharge power capacity [F] |  |  |  | 0.0071
|  | Ƭ Theorique |  | Ƭ=RC
Nom du test |  | FINAL10
Final
Tests effectués le : |  | 2013-01-07
Nom de l'opérateur : |  | FOP
Tamb [°C] |  | 20.3
Humidité [%] |  | 27.1
Constante de temps Ƭ |  | Ƭ Mesurée |  | Ƭ Theorique
|  | [ms] |  | [ms]
QH7L |  |  |  | 52.54
QH7R |  |  |  | 52.54
QH3L |  |  |  | 52.54
QH3R |  |  |  | 52.54
QH8L |  |  |  | 52.54
QH8R |  |  |  | 52.54
QH9L |  | HS - different response |  | 52.54
QH9R |  |  |  | 52.54
Graphique